## Slide 1: Slow Extraction workshop Absolute calibration of the secondary emission monitors at CERN J. Bernhard, M. Brugger, N. Charitonidis, M. van Dijk, A. Goillot, M. Fraser, N. Menaa, M. Munos, L. Parsons França, F. Ravotti, F. Roncarolo, P. Simon, J. Tan.

## Slide 2: The intensity monitors of the North and East Area
The CERN North and East Area serve a wide variety of fixed target experiments and testbeam users with tailor-made primary and secondary beams
  For very high-intensity beams (1011-1013 protons) there are not so many detectors capable of measuring the absolute intensity
  Choice has been SEM-based detectors
SEM (Ti Plate)
e-
P+
BIAS
BIAS
Signal
Number of protons in North and East Area primary lines monitored with SEM detectors
  It is a beam intercepting device, collecting the small fraction of low-energy secondary electrons emitted from a metal foil when the beam passes through: small current for large beam
  Due to high intensity (1012-1013 protons at 400 GeV/c) but slow extraction (4.8s), other detectors are not considered viable for absolute intensity monitoring
F63.XSEC008 during maintenance
F61.XSEC023 during maintenance
Photos from: A. Goldblatt
L. Parsons Franca

## Slide 3: SEM detectors in North Area
210279
210216
211626
220412
251010
251247
251248
240610
241149
241150
230949
230950
230705
043282
045769
045914
SPS
The TT20 primary line extraction protons from SPS to the North Area is covered with SEM detectors such that at all points the intensity can be reasonably known
  Extraction to North Area from SPS features two splitters (S1, S2)
  Two SEMs in the upstream section
  SEM before and after each splitter
  One SEM upstream (40-50m) before each target
  Two SEMs in each target station (one in T10)
    TBIU (Target Beam Instrumentation Upstream)
    TBID (Target Beam Instrumentation Downstream)

## Slide 4: SEM detectors and the difficulty of absolute intensity
Long-standing uncertainty on NA intensity
  Last calibration >20 years ago (?)
  Expected ~10%, to have a few % would be good
Observed since many years
  SEM intensity != SPS BCT intensity
  Unstable over year and from year to year
  Not clear what part is the monitor and what part is losses
  Displayed most upstream monitor in TT20 (210279, two foils)
Gathering experience in diagnostics
  Analyze data from all SEM monitors in North Area
  Compare monitors with each other
  Develop the diagnostics that will lead to a strategy for addressing the issues
 | Slope foil A | Slope foil B
2021 | 0.569 +/- 0.005 | 0.690 +/- 0.004
2022 | 0.634 +/- 0.003 | 0.731 +/- 0.003
2023 | 0.736 +/- 0.005 | 0.814 +/- 0.005
Foil A
Foil B
Preliminary
Preliminary
Preliminary

## Slide 5: Damage to SEM foils
Year to year, the sensitivity of the foils seems to change
  Likely caused by small differences in hit position
  Each part of the foil has a different history
Beam movement on target monitor (TBIU)
  Changing secondary momentum horizontally displaces beam on SEM
  Movement on SEM monitor can be ~2mm
  Spot size on T4 has σx ~0.4mm (T2 likely similar)
The SEMs in the target areas are likely not giving an accurate assessment of intensity
  Foils in TBIU more damaged than upstream due to small spot size (damage scales with area)
  Estimated intensities could be wrong by >20%
T2 target TBIU (2022)
T10 target TBIU scan
Preliminary
Preliminary

## Slide 6: Calibration technique – activation foils
Foils placed in beam are activated mostly by fission driven by the proton beam
  Use cross-sections from literature to assess activation
  Cross-section values are not unambiguous: the reference type and value matters
    For example, the values for isotope production cross-sections from natural copper are usually normalized with respect to values of 24Na measured in a reference aluminium foil
    Across different publications, the value used for normalization differs, and the error of the reference is not included
    Assess appropriate value, homogenize cross-sections and account for reference error where needed
Foils assessed by means of gamma spectrometry
  Correct for half-life of isotopes
  Factor in cooldown time between irradiation and correct for growth / decay over irradiation

## Slide 7: Calibration of SEMs in East Area
SEM-based intensity monitors
  Initial comparison between fast BCT and foil activity (2022)
    Measurement based on single isotope (24Na in aluminium foils)
    Compared with fast extracted beam (few 10s of ns) onto fast BCT
    Methods agree with absolute difference of 9.3%
  Second irradiation with slow extracted beam (2022)
    Measure SEM signal & compare it with foils activity
    SEM signal calibration based on resulting measurement
  Data collected also with copper foils in 2023
    Analysis ongoing and promising
Intensity calibration of T9/T10 primary lines not yet performed
  Possibly the factors of T8 can be reused – similar / same detectors
  Will be assessed further once analysis of T8 foils is completed
F. Ravotti, EP-Tech-Note-2022-001

## Slide 8: Calibration of SEMs in North Area
Valuable experience gained in 2022 and 2023
  Put aluminium / copper foils in line (in this case, near target)
    Aluminium is “standard candle”, copper has many isotopes -> statistics
  Expose to beam (~100-200 shots), sum up SEM signals
  Cross-sections gathered for proton beam >= 300 GeV/c
  Divide activity by cross-section, and compare to SEM sum
First set of calibrations done with T10 target
Second set of calibrations done for T2+T4+T6
  Measured fewer POT in activation foil than on SEM
Third set of foils measured in HiRadMat (later slide)
Aluminium foils give different results than copper
  Impact of neutrons through auxiliary process
  Present in both TCC2 and HiRadMat measurements
Still preliminary
Ratio of activation and SEM
Still preliminary

## Slide 9: Cross-section measurement in
Series of foils inserted in HiRadMat facility at CERN
  Foil stack of 3x100µm pure aluminium, 3x100µm pure copper
  440 GeV/c fast-extracted beam, accumulated 9.078e13 POT (ring BCT, assume 2% error)
  Substantial material on beam axis (BTV, 6mm carbon) and 3.5m air
Results shown for isotopes measured in copper compared to literature values
  Literature values for “low” energy (28-30 GeV) and high-energy (150-800 GeV/c)
  All sources using 27Al(p,3pn)24Na as reference normalized to 8.35 +/- 0.4mb (for all energies)
  Fit to data gives slope 0.991+/- 0.015 (low energy) Fit to data gives slope 1.000+/- 0.014 (high energy)
  Cross-sections not sensitive to energy (28-800 GeV/c)
  No measured cross-section increase due to material (64Cu not shown, see next slide)
47Ca
62Zn
48Cr
Isotopes measured in copper foil
28Mg
48Sc
57Ni
55Co
43K
59Fe
47Sc
42K
24Na
56Co
44mSc
52Mn
46Sc
48V
7Be
54Mn
51Cr
57Co
58Co
Cross-section from measurement (mb)
Still preliminary

## Slide 10: Cross-section measurement in
Both aluminium and copper foils indicate a substantial neutron flux (from material on the beam axis)
  Modifies production of 64Cu in copper 63Cu(n,γ)64Cu
  Modifies production of 24Na in aluminium 27Al(n,α)24Na
Material correction for aluminium (1.5% +/- 0.5%) / 100mg/cm2
  1520 mg/cm2 (for aluminium foil) leads to 8.15 +/- 0.48 mb for 24Na
  Correction not well-established for copper foils (and not so relevant)
Results from HiRadMat will be used to update North Area SEM calibrations – substantial effect
 | Literature CS | Observed CS
24Na from Al foil | Between 7.9-8.8 mb | 9.86 +/- 0.32 mb
64Cu from Cu foil | 14.2 +/- 0.7 mb | 23.4 +/- 2.7 mb
EXFOR data
EXFOR data
Still preliminary

## Slide 11: Intensity of DC beams in North and East – options
For future: three possible technologies under consideration: SEM, BCT and CCC
Secondary Emission Monitors are the currently implemented technology
  SEM foils in North Area seem to get substantially damaged by beam – possibly even over a single run-year
  Coupled to movement of the beam and vacuum history this gives a large (but slow) variations
Beam Current Transformers have been considered (but mostly ruled out)
  Ideally, measurement is “short” (1ms or less) – but even fast-pulsed slow extraction to TT20 is 10-20ms
  Current lowest current is ~200µA in a ~200µs pulse (end of LINAC3)
  Regular slow extraction is 4.8s (equivalent to 1.3µA with 4 1013 protons) – unfeasible
  Fast extraction (using kickers) is limited in intensity allowed to be extracted to North Area
Cryogenic Current Comparator is an excellent candidate (see next talk!)
  Magnetic field of beam induces screening current in superconducting shielding
  A single CCC could be used as an absolute standard for the existing intensity monitors
    Should be located near a SEM to optimize further cross-instrument studies
  Technically feasible, being investigated for deployment in Run4 (2028-2029)
J. Tan

## Slide 12: Next steps & future plans
Observation: Continuous degradation of SEM monitors
  Consequence: uncertainty of 20% up to 50%, will grow in the future
  Full impact on beam operation to be assessed
  Will finish analysis of available data far from fully analyzed and condense into strategy
Plan for the future
  SEMs are difficult, BCT not (very) possible, but CCC for Run4 (2028-2029) is technically feasible
    Finally, CCC is the only viable choice for a sustainable solution
  Beam time for two “full” calibrations of North Area requested for 2024
    Continue calibration efforts with activation foils until full consensus found
    Cross-calibration of SEMs and CCC could be key
  Fast foil exchange is crucial for the activation foil calibrations
    Investigate mitigation: motorize SEMs? (stepping motors instead of in/out)
    Investigate mitigation: adapt SEM design for easier calibration?

## Slide 13: Conclusions
SEM detectors based on the principle of secondary emission are used for monitoring the intensity in the CERN East and North area
  Extensive fluctuation over time of the detector has been shown
  Substantial damage to the foils due to extensive exposure of the beam has been found
Activation foils are used to assess the absolute intensity of the beam for recalibration
  Last done >twenty years ago in the North Area, experience is being rebuilt
  Experimental results from the HiRadMat facility will be used as the foundation for further calibrations
  Not fully clear if aluminium is suitable as the baseline foil due to neutron sensitivity
For the future, multiple different strategies will converge
  Continue building expertise and calibrate SEM-based detectors with activation foils as possible
  Investigate (one) Cryogenic Current Comparator for an online absolute calibration standard (2028-2029)
    Should be located near a SEM to optimize further cross-instrument studies

## Slide 14: Closing thoughts
Many thanks to the people who contributed to this talk!
J. Bernhard, M. Brugger, N. Charitonidis, A. Goillot, M. Fraser, N. Menaa,
M. Munos, L. Parsons França, F. Ravotti, F. Roncarolo, P. Simon, J. Tan.

## Slide 15: Backup slides

## Slide 16: SEMs at the TCC2 targets (T2, T4, T6, T10)
251010
251247
251248
240610
241149
241150
230949
230950
230705
045769
045914
Each target station has three SEM monitors associated
    1x “Upstream”
    2x target instrumentation (just before and after)
      T10 has no TBID
The estimated beam intensity comes from the “upstream” SEM
  230705(T2), 240610 (T4), 251010 (T6)
  For T10, the estimated intensity comes from the TBIU (045914)
If all goes well, the ratio TBIU / upstream = 1
    Upstream monitors are only ~40-50m upstream

## Slide 17: Monitor BSI.210279 (two foils)
SPS ring
TT20
Monitor BSI.210279
Monitor had some issue over part of 2023
Preliminary

## Slide 18: Monitor BSI.210279 (two foils)
Ratio of two foils in the same location
Preliminary

## Slide 19: Comparison TBIU and upstream SEM – T2, T4, T6, T10
Use ratio (TBIU / upstream SEM) as an indicator of stability
  Indication of good calibration: value equal to 1
  Indication of stability: line is flat
Upstream monitor of T10 had some issue in 2023 (?)
Preliminary

## Slide 20: Layout of the TBIU
It was observed at the T10 target that the foils get damaged
  Scales with size of beam – more protons per mm2
TBIU has two “heads”, A and B
  Beam steering (symmetry) happens on two BSPs
  Both of these are located in head A
  Assume beam is centered perfectly on those
  However, the SEM is located in head B
  Distance between the two is around 12.5cm
BSPs
SEM

## Slide 21: Movement of the beam (example at T4)
Wobbling magnets near target allow to angle beam onto target, selecting different momenta for beamlines
Different wobbling settings give a different angle on the target
    The TBIU and TBID stations move with the wobbling setting
    The beam is then steered using the BSP settings
    However, lever arm of 12.5cm means beam is displaced on the intensity monitor
    Total difference (for two calculated wobblings) is ~15.7 mrad
    Absolute displacement of almost 2mm
Combine with beam size at T4 target
    In short, beam moves five sigma between these two settings
T4 standard wobbling
H8 (+180), H6 (+120), P42 (+400)
H8 (-300), H6 (-120), P42 (+400)
Primary beam size T4 (G. Mazzola)

## Slide 22: Copper cross-section comparison (linear scale)
Linear version of plot shown on earlier slide
  Comparison of absolute cross-sections measured in copper foil exposed to beam in HiRadMat facility
Preliminary
Preliminary