## Slide 1
Slow Extraction Workshop, February 11th – 15th, MedAustron
Thomas Sieber (GSI)
for the CCC Collaboration
The Cryogenic Current Comparator (CCC) for nondestructive Spill Measurement in the nA range
Notes: 1

## Slide 2
Outline
CCCs at FAIR, Motivation and working Principle
Laboratory Tests and Operation in CRYRING@ESR
Measurement Campaign 2023
Alternative CCC Types, Pros and Cons
Summary and Outlook
Slow Extraction Workshop
Thomas Sieber 2
February 13th 2024
Notes: 2

## Slide 3
GSI
FAIR
SIS100
HESR
Cryring
SIS18
UNILAC
Proton Linac
CR
CBM
Super FRS
NuSTAR
PANDA
CCC Position
CCC:
Precise measurement of
azimuthal magnetic field
(fT-range) using
DC-SQUIDS
Low Intensity Measurement at FAIR
Nondestructive measurement for low intensity
(nA) beams at FAIR
- Slow extracted beam from SIS100 (RIB, CBM), spill optimization
- Circulating beam in Collector Ring (CR  HESR)
Detector systems used for slow extraction at SIS18 (P. Forck)
SClNT …Scintillator
IC …Ionization Chamber
SEM …Secondary Electron Monitor
Slow Extraction Workshop
Thomas Sieber 3
February 13th 2024
Notes: 3

## Slide 4
CCC Working Principle
 commercial
products (MAGNICON, SUPRACON)
+
- FLL electronics:  Resolution
Proof of Principle in Accelerators: A. Peters at. al. 1994
Slow Extraction Workshop
Thomas Sieber 4
February 13th 2024
Notes: 4

## Slide 5
Cryostat Design for CRYRING / FAIR
System costs (2019)
Nb CCC 180 k€
Cryostat 150 k€
Liquefier 65 k€
SQUID + electronics 10 k€
DAQ 10 k€
Support 5 k€
Total ~ 420 k€
250 mm
60 kg
1200
Liquefier
Head
Vacuum Tank
Aluminum
Windows
Thermal
Shielding
with
Enthalpie Line
UHV Beam
Tube
Helium
Container
CCC Shield +
Detector
Steel
Frame
Cryostat Design
Slow Extraction Workshop
Thomas Sieber 5
February 13th 2024
CCC Installation
in CRYRING
standing time ~ 10 days
Notes: 5

## Slide 6
Laboratory Tests I
10 kHz Filter ON
direct
balanced
Pulsed Current Measurement from Calibration Loop
Direct = without Matching Transformer / Balanced = Matching Transformer adjusted
balanced
10 kHz Filter OFF
noise spectrum
Slow Extraction Workshop
Thomas Sieber 6
February 13th 2024
Matching Transformer
Notes: 6

## Slide 7
Laboratory Tests II
FLL-Electronics
hard
soft
direct: 260 nA @ 200 kHz
=> 0.33 µA/µs
balanced: 130 nA @ 200 kHz
=> 0.16 µA/µs
Tolerable rise time of pulses in flux locked mode  Slew Rate
Slow Extraction Workshop
Thomas Sieber 7
February 13th 2024
Notes: 7

## Slide 8
Low beam currents
Resolution limited by external perturbations (liquefier, dipole ramp)
hardware optimization
improve software filtering
Operation in CRYRING: Challenges at low and high Intensities
High beam currents
High slew rates  loss of signal lock
Threshold around 500 nA of D+
add low pass to reduce slew rate
(10 kHz instead of 100 kHz cut off)
Slow Extraction Workshop
Thomas Sieber 8
February 13th 2024
In general: transition laboratory  accelerator reduces current resolution by factor 10

## Slide 9
Results from CRYRING
Low Intensity
High Intensity
Slow Extraction Workshop
Thomas Sieber
February 13th 2024
9

## Slide 10
Spill Measurements at GSI - History
Spill measurement with CCC and SEM
FFT
Spill analysis unbunched beam
Bunched beam
Kurian et. al. 2015, Ni26+, 600 MeV/u, textr 60 ms – 5 s; Magnicon SQUID + electronics
Slow Extraction Workshop
Thomas Sieber 10
February 13th 2024
Experimental setup
CCC prototype (A. Peters et al.)
Forck et. al. 2004, Ne10+, 300 MeV/u, textr 100 ms – 5 s; FSU Jena SQUID + electronics
Change requests for FAIR:
Beamline and isolation vacuum coupled
~5*10-6 mbar  no UHV operation
Shield not completely enthalpy cooled
additional cold head required
Cooler to be switched off during
measurements (thermal drift)
No liquefier, manual filling  (no standalone operation AND thermal drift)
SQUID, electronics + controls from 80‘s
Magn. shielding, current resolution etc.

## Slide 11
Slow Extraction Workshop
Thomas Sieber 11
February 13th 2024
FFT spectrum
Engineering Run 12/2023 CCC-XD@SIS18
Crescimbeni et. al. 2023, U73+, 200 MeV/u, textr 50 ms – 1 s; Magnicon SQUID + electronics
Experimental Setup
CCC installed in exp. cave

## Slide 12
Slow Extraction Workshop
Thomas Sieber 12
February 13th 2024
Engineering Run 12/2023 CCC-XD@SIS18
Eliminating Pertubations, investigation of filtering effects (spill shape, particle numbers) …
liquefier off
sine wave superimposed
liquefier on, long spill
liquefier on, short spill
liquefier off, short spill
+ calibration signal
(ms)

## Slide 13
Engineering Run 12/2023 CCC-XD@SIS18
Slow Extraction Workshop
Thomas Sieber 13
February 13th 2024
Analysis of spill optimization, tune wobbling …
… and bunching
500 Hz
3 kHz
9 kHz
500 Hz
9 kHz
3 kHz
0 kHz
bunched
unbunched
I (µA)
I (µA)
t (ms)
t (ms)
bunched
wobbled
FFT spectra for different wobbling frequencies

## Slide 14
Alternative CCC Types
Dual Core CCC
Slow Extraction Workshop
Thomas Sieber 14
February 13th 2024
Noise figures for classical- axial- and dual-CCC
Axial/Coreless-CCC
Pros
magn. shielding (150 dB instead of 75 dB)
lead / costs
easy manufacturing
[two stage SQUID]
Cons
low coupling
high noise background
Dual-CCC
Pros
magn. shielding (150 dB instead of 75 dB)
lead / costs
easy manufacturing
two stage SQUID
strong coupling
low noise background
redundancy

## Slide 15
Summary and Outlook
CCC focus at FAIR changed from storage rings to transport sections  FAIR Early Science … FS, FS+ scenarios
CCC-XD successfully operated in storage rings and transport lines, nA and spill measurement could be demonstrated, system fullfills FAIR requirements (almost)
Beamtime in Dec. 2023showed capability for online spill analysis
Axial CCC type shows excellent magnetic shielding, but noise sensitivity too large
Dual core CCC combines advantages of both CCC systems and will be our choice for FAIR
DCCC recently tested in Jena laboratory and ready for installation. Beamtime in April
Todos:
- New thermal shield test (Mai 2024)
- Liquefier test (June 2024)
FESA class and slow controls (cryo system) development ongoing, including all filtering, online data processing and SQUID controls (interaction with spill optimization)
Calibration pulse measurement with Dual CCC, Feb. 2024
Summary and Outlook
Slow Extraction Workshop
Thomas Sieber 15
February 13th 2024

## Slide 16
CCC Collaboration 2015 - 2024
HI Jena: T. Stöhlker, V. Tympel
FSU Jena: [R. Neubert, P. Seidel] T. Köhler, J. Golm, F. Schmidl
GSI: L. Crescimbeni [D. Haider] T. Sieber, M. Schwickert
HIT Heidelberg: A. Peters
CERN: [M. Fernandes] J. Tan, T. Koettig
IPHT: R. Stolz, M. Schmelz, V. Zakosarenko
TEMF: H. De Gersem [N. Marsic] W. Müller
Slow Extraction Workshop
Thomas Sieber 16
February 13th 2024
Notes: 16

## Slide 17
Thank you for your attention !
Notes: 17

## Slide 18
Appendix
8. Annual MT Meeting, DESY, Hamburg
Thomas Sieber
September 27th 2022
Notes: 18

## Slide 19
SQUID principle uses two quantum effects:
Josephson Effect Flux Quantisation
CCC Working Principle II
Courtesy R. L. Fagaly, Tristan Techn.
Matching Transformer
SQUID Cartridge
Schematic of
the FLL /SQUID
electronics
Slow Extraction Workshop
Thomas Sieber 5
February 13th 2024
Notes: 19

## Slide 20
CCC Working Principle II
Evolution from cylindrical over coaxial tp meandrical shielding
Analytical approach by PTB:
Damping is function of radius ratio (Rout/Rin) and number of meanders
COMSOL + Microwave- Studio calculations for field penetration into the shielding
Example Dipol field:
Attenuation achieved:
~150 dB
BUT:
Beam field (10nA): 50 fT
Earth magn. field: 50 µT
to SQUID
8th Annual MT Meeting, DESY, Hamburg
Thomas Sieber
September 27th 2022
Notes: 20

## Slide 21
Production of the CRYRING Cryostat
Frame
He-Chamber
Thermal
Shield
He-Chamber
with MLI
Shield and
tubing with MLI
Al Windows
mounted
8th Annual MT Meeting, DESY, Hamburg
Thomas Sieber
September 27th 2022
Notes: 21

## Slide 22
Temperature / Pressure Offset Drift
Offset temperature/pressure dependence
Baseline drift of ~1nA/s observed during beam tests, curve flattens when thermal equilibrum is reached
If T/P effect, high relevance due to backpressure +
fluctuations from recovery lines
offline tests with closed exhaust line
- Effect investigated at Jena with 'anticryostat' setup,
decoupling T and P
Reason identified as T-dependent inductance variation
 can this be avoided ?  coreless CCC ?
GSI Accelerator Salon
Thomas Sieber
26th February 2020
Notes: 22
22

## Slide 23
Signal from helium
refrigerator (1.4 Hz)
Dipole ramp
deterministic:
subtracted using transfer function
calculated from
background measurements
x
=
dipole field
transfer function
correction
(no mechanical coupling!)
Beam Operation: Pertubations and Filters
Main disturbances: liquefier pulse + dipole ramp
(Origin)
Slow Extraction Workshop
Thomas Sieber 10
February 13th 2024

## Slide 24
Investigation of new Shielding Design
Simulations performed with MWS to
check shielding properties
Assumptions:
- Material = perfect diamagnet
- Shield currents + toroid not included
- Homogeneous dipole field used
- Real meander geometry used
'classic'
'outer'
'inner'
classic
Shielding design for the FAIR and AD CCC
GSI
AD
FAIR
Attenuation decreases for larger Rin
For given Rin: Attenuation increases for larger Rout
Rout related to total shield mass
and cryostat dimensions
Number of meanders increases attenuation
CCC-Prototype: Rin=75 mm, Rout=125 mm  Att. 108 dB
FAIR-CCC: Rin= 125 mm, Rout=175 mm  Att. ~ 80 dB
 4 more meanders required for comparable attenuation
HIJ Semiannual Palaver
Thomas Sieber
6th October 2021
Notes: 24

## Slide 25
GSI Accelerator Salon
Thomas Sieber
26th February 2020
Laboratory Tests: Comparison of CCC Types
Notes: 25

## Slide 26
SQUID Development
GSI Accelerator Salon
Thomas Sieber
26th February 2020
IPHT cross type JJ technology
sub µm thickness
low junction capacitance
high voltage swings without
hysteresis,
 low noise
high tolerable background fields
inductance
Notes: 26