## Slide 1
Measuring Small AC Ripple on Strong DC Currents
Frank Stulle, Laurent Dupuy, Hervé Bayle (Bergoz Instrumentation)
Slow Extraction Workshop 2024-02-11 - 15

## Slide 2: Motivation
D. Prokopovich et al., “Slow Extraction at MedAustron”, I.FAST REX Kick-Off Meeting, Feb 2021
Stability of slowly extracted beams ⇕ Stability of magnet power supplies ⇕ Quality of feedback regulation ⇕ Precision of power supply current measurement
Beam

## Slide 3: Development Objective
I.FAST Work Package 5, Task 5.3: “Improvement of resonant slow extraction spill quality”
I.FAST proposal: measure magnet currents to
Later communication: measure magnet currents to

## Slide 4: Short Interlude
Rewind to Monday 2024-02-12 16:30
Ryotaro Muto, “Spill structure with newly upgraded main magnet power supplies in J-PARC Main Ring”

## Slide 5: Short Interlude
Question by audience: “Does your DCCT really have 10-7 resolution?”

## Slide 6: Short Interlude
My answer: “Be careful!”
Spectral amplitudes may be at such low levels.
But to get resolution we need to apply the inverse Fourier Transform.
Since spectrum has a certain bandwidth, what we get in time domain is a signal averaged over a certain time window.
Trade-off: good amplitude resolution vs good time resolution!
To my knowledge, there is no current measurement system available that achieves close to 10-7 resolution with >100 Hz bandwidth.

## Slide 7: Short Interlude
Fast Forward to Wednesday 2024-02-14 14:00

## Slide 8: Development Objective
Measurement bandwidth to
Magnet currents up to
Current modulations up to
Collaboration consists of several groups with very different accelerators and very different magnets.
That means, any development must be adaptable to very different requirements.

## Slide 9: Development Starting Point
Required AC resolution relative to : (⇒ 140dB dynamic range)
Required AC resolution relative to : (⇒ 100dB dynamic range)
Required DC accuracy and resolution not properly specified, but certainly a lot less stringent than 10-7.
⇒ It seems advantageous to separate DC and AC measurements.

## Slide 10: Development Starting Point
We are current transformer specialists.
Hence, we decided to base the development on a current transformer.
AC resolution and bandwidth are almost achieved by ACCT:
ACCT bandwidth 1 Hz – 1 MHz
ACCT resolution <5×10-5 relative to full scale current
Assuming flat noise spectrum, low pass filtering to 40000 Hz should result in 1×10-5 resolution.

## Slide 11: Challenge
Presence of strong DC currents, i.e. strong magnetic fields ⇒ saturation of magnetic materials used in current transformers
A few Amperes suffice to saturate high performance magnetic materials.
Need to wrap the current transformer in a DC compensation coil.
Required compensation accuracy: ≲1 A

## Slide 12: Basic Principle

## Slide 13: Basic Principle
Design choice: both output signals shall be properly in phase.
That means, to reconstruct the full input current spectrum (DC – 40000 Hz):
(DC – 10 Hz)
(10 Hz – 40000 Hz)

## Slide 14: Prototype Specifications
DC: 8 V for 6000 A, DC – ~10 Hz
AC: 10 V for 60 A, ~10 Hz – ~40000 Hz
Two windings, each 1500 turns
110 mm aperture
260 mm outer diameter
80 mm height

## Slide 15: Prototype Tests at CERN
6000 A DC Test Stand
Reference Fluxgate DCCT
Prototype Sensor
Reference Fluxgate DCCT Electronics
Prototype Electronics
Precision Voltmeters

## Slide 16: DC Ramp Up and Down
Prototype offset and gain corrected to match reference DCCT
Curves almost exactly superimpose

## Slide 17: Prototype Tests at BI
Power Amplifier with Load
Prototype Sensor 60 turns primary
24bit ADC
Prototype Electronics
Signal Generator
Feedback Power Supply

## Slide 18: Frequency Response

## Slide 19: DC + Small AC
500 ADC + 0.5 App 70 Hz sine
Curves almost exactly superimpose

## Slide 20: Measurements Close-Up
well?
What went
wrong?

## Slide 21: Close-Up DC
stable during flat top
too high error during ramp

## Slide 22: Close-Up Frequency Response
Measured frequency response not sufficiently flat at very low frequencies (<10 mHz) (not visible in simulations)

## Slide 23: Close-Up DC / Frequency Response
Apparently, DC compensation was behaving slightly wrong during measurements at CERN.
Some shortcomings were quickly found, others were well hidden.
Nothing was fundamentally wrong.
Re-designed electronics for improved signal accuracy and stability.
New electronics is being assembled now and will be tested soon.

## Slide 24: Close-Up AC
Noise = 2.6 mArms ⇒ 4.3 × 10-5

## Slide 25: Close-Up AC
Input Current: 500 ADC + 0.5 App 70 Hz sine
AC Noise:
DC Feedback=Off , Input Current: no DC, ≥0 AAC
AC Noise:
Spec: , 40 kHz BW
same with 0 A input current

## Slide 26: Summary
High resolution, high dynamic range current measurement system in development for stabilization of magnet power supplies.
Prototype measurements were rather satisfactory.
Bandwidth surpasses specifications.
Measured AC resolution is (almost) within specifications. And there seems to be potential for further improvements.
However, DC compensation worked less well than expected.
Some corrections and improvements have been implemented already.
New electronics will be finished this week.