## Slide 1: Slow Extraction Developments at MedAustron
Florian Kühteubl 5th Slow Extraction Workshop 14.02.2024

## Slide 2: Slow Extraction at MedAustron
Beam accelerated to target energy
Phase Jump to increase momentum spread
Resonant sextupole ramped to prepare resonance
Extraction optics
Extraction optics 𝑄_𝑋=5/3
𝑄_𝑥^′=−4 Δ𝑝/𝑝=−2.3𝐸−3
Betatron Core accelerates beam into resonance
Beam accelerated to target energy
Phase Jump to increase momentum spread
Resonant sextupole ramped to prepare resonance
Betatron Core accelerates beam into resonance

## Slide 3: Steinbach Diagram
~ΔQ
= 1.6666
Unstable area
Stable area
Beam

## Slide 4: BETATRON CORE EXTRACTION
Betatron Core accelerates particles
Particle momentum is increased
Beam is pushed into resonance

## Slide 5: RADIO FREQUENCY KNOCK OUT (RFKO) EXTRACTION
Particle amplitude is increased
Transverse kicker with AC
Frequency matched to beam revolution

## Slide 6: ADVANTAGES OF RFKO
Compatible with bunched beam extraction
Sub-millisecond adaptation of the beam intensity
 Dynamic Intensity Control
 Irradiation Gating
Compatible with FLASH treatment
Ripple reduction by adjusting the excitation pattern and introducing a closed-loop feedback regulation
Changing the beam energy during the spill  Multi Energy Extraction

## Slide 7: RFKO SIMULATION - PROOF OF PRINCIPLE

## Slide 8: RFKO PARAMETER OPTIMIZATION
The RFKO extraction settings can be divided in 2 categories:
The machine parameters:
  Tunes, Chromaticity, , phase-advances,…
  Impact the extracted beam condition at ESE: spiral steps (beam size), beam loss at ESE, angle spread, energy spread, transmission
RFKO Signal parameters:
  The extraction methods, Kick amplitude, RFKO frequency spread, central frequency, Time sweep (for FM),time randomness and phase offsets (for PSK)
   Impact the extracted spill condition (smoothness, spill length, ripples…)
ESE
ESE
Courtesy of Fabien Plassard

## Slide 9: Machine Parameter optimisation
Extraction efficiency
Transmission efficiency
Beam distribution at septa
Dispersion

## Slide 10: RFKO Signal Optimisation
Frequency
modulation
Pseudo Random Binary Phase Shift Keying

## Slide 11: RFKO SETUP
Ettus X310 SDR
Courtesy of X. German

## Slide 12: MEASUREMENT RESULTS
Courtesy of D.A. Prokopovich
First RFKO extraction at MedAustron!

## Slide 13: RFKO Beam Quality
For Comparison: Betatron Core w/ EBC: RF = 3
Betatron Core w/o EBC: RF = 12

## Slide 14: CONSTANT OPTICS SLOW EXTRACTION (COSE)
Bunched operation
Multi Energy Extraction
Resonance is moved into the beam
Linear ramp of all MR magnets
Synchronized ramp of radial beam position

## Slide 15: MULTI ENERGY EXTRACTION WITH COSE
En 1
En 2
En 3
En 4
En 5
En 6
En 7
En 8
En 9
En 10
Ten energies with

## Slide 16: PHASE DISPLACEMENT EXTRACTION (PDE)
Fast extraction times
FLASH
Empty bucket is moved through the waiting beam
Sweep can be repeated multiple times

## Slide 17: PDE measurements

## Slide 18: PDE Measurements
Amplitude modulation
Courtesy of E. Renner

## Slide 19: Ripple Reduction
Less ripples
Constant dose rate
Faster treatment
Less intensity fluctuations

## Slide 20: SUMMARY
Successful testing of alternative extraction methods (RFKO, COSE and PDE)
Proof of principle: extraction of the circulating beam into the treatment room
Further optimization and ongoing developments needed
Understanding of complex beam dynamics during extraction

## Slide 21: THANK YOU
⚛ Xavier German
⚛ Fabien Plassard
⚛ Dale Prokopovich
⚛ Elisabeth Renner
⚛ RFKO Team: ⚛ Thomas Margreiter
⚛ Laurids Adler
⚛ Felix Feichtinger
⚛ Lorenz Fischl
⚛ Greta Guidoboni
⚛ Katrin Holzfeind
⚛ Clemens Maderböck
⚛ Mauro Pivi
⚛ Alexander Wastl
⚛ Valeria Rizzoglio
⚛ Claus Schmitzer
⚛ Ivan Strasik
⚛ Michael Benedikt
⚛ Rebecca Taylor
⚛ Pablo Arrutia Sota
⚛ Matthew Fraser

## Slide 22: THANK YOU FOR YOUR ATTENTION!