## Slide 1
GOVERNMENT OF ROMANIA
Structural Instruments
2007-2013
Sectoral Operational Programme
„Increase of Economic Competitiveness”
“Investments for Your Future”
Extreme Light Infrastructure – Nuclear Physics (ELI-NP) - Phase I Project co-financed by the European Regional Development Fund
Preparation for future photoneutron experiments at ELI-NP
Dan Filipescu
EUNPC, Groningen
31 August – 4 September 2015
EUROPEAN UNION

## Slide 2: Outline
Future experiments at ELI-NP – Gamma Above Neutron Threshold
Discrepancies between existing photoneutron cross section measurements
Experimental challenges for neutron multiplicity sorting
Design of a flat-efficiency neutron detector
Experiment: 209Bi(g,xn) cross sections at NewSUBARU synchrotron radiation facility

## Slide 3
Laser Compton Scattering (LCS) gamma ray sources
Main concepts

## Slide 4: ELI-NP -ray beam expected parameters

## Slide 5: ELI-GANT – Physics program

## Slide 6: ELI-GANT – Physics program

## Slide 7: Purpose of present study - Measurements of partial and total photoneutron cross sections for 209Bi
Most of the photoneutron cross section measurements were performed using quasi-monochromatic annihilation – QMA photons using positron in flight annihilation at two major facilities:
  Saclay (France)
  Lawrence Livermore National Laboratory (USA)
Large discrepancies in (,xn) c.s. measured at the two facilities:
  (γ, 1n) c.s. are generally noticeably larger at Saclay than at Livermore
  (γ, 2n) c.s. are generally larger at Livermore than at Saclay.
Livermore - 84
Saclay - 80
Both – 42
Main problem
Many bremsstrahlung
data
Other - 10
No systematic way to resolve the discrepancies:
New and reliable measurements are required!

## Slide 8
For (,1n) cross sections – measurements using LCS -ray beam and the high efficiency neutron detector developed at GACKO offer the required precision and reliability
E<S2n  (,1n)
Average energy Eavg determined using the ring ratio method

## Slide 9: For (,1n) cross section measurements
Average neutron energy and neutron detection efficiency determined using the ring ratio technique

## Slide 10
Example:
E<S3n  (,1n) + (,2n)
Neutrons from (g,1n) reactions
Average energy E1
Energy range: 0 to Eg-Sn
Neutrons from (g,2n) reactions emitted by the:
AX nucleus
  Energy range: ~( 0 to Eg‒S2n)
A-1X nucleus
  Energy range: ~( 0 to Eg‒S2n)
Average energy E2
E1  E2  ε(E1)  ε(E2)
For (,2n), (,3n), … cross sections

## Slide 11
Let us assume E<S4n  (,1n), (,2n), (,3n) induced reactions
N1, N2 , N3 = number of (g,1n), (g,2n), (g,3n) induced reactions
E1, E2 , E3 = average energy of (g,1n), (g,2n), (g,3n) neutrons
Ns, Nd , Nt = number of single, double and triple neutron events
Adjust experimental conditions to have maximum one reaction per -ray beam pulse
Ns, Nd , Nt - we measure directly
N1, N2 , N3 - we must determine to obtain the cross sections
E1, E2 , E3 - ? Can not be determined using the ring ratio method PROBLEM
A neutron multiplicity sorting technique with a flat efficiency neutron detector is required!

## Slide 12
Let us assume E<S4n  (,1n), (,2n), (,3n) induced reactions
N1, N2 , N3 = number of (g,1n), (g,2n), (g,3n) induced reactions
E1, E2 , E3 = average energy of (g,1n), (g,2n), (g,3n) neutrons
Ns, Nd , Nt = number of single, double and triple neutron events
Adjust experimental conditions to have maximum one reaction per -ray beam pulse
Ns, Nd , Nt - we measure directly
N1, N2 , N3 - we must determine to obtain the cross sections
E1, E2 , E3 - ? Can not be determined using the ring ratio method PROBLEM
A neutron multiplicity sorting technique with a flat efficiency neutron detector is required!
Single neutron events:
Double neutron events:
Triple neutron events:

## Slide 13
Considering a flat efficiency neutron detector and E<S4n we have:
Ns, Nd , Nt - we measure directly
ε - known by direct measurement and by simulation
N1, N2 , N3 - we determine to obtain the (,1n), (,2n), (,3n) cross sections

## Slide 14
Considering a flat efficiency neutron detector and E<S4n we have:
Single neutron events:
Double neutron events:
Triple neutron events:
Ns, Nd , Nt - we measure directly
ε - known by direct measurement and by simulation
N1, N2 , N3 - we determine to obtain the (,1n), (,2n), (,3n) cross sections

## Slide 15
Neutron multiplicity sorting with a 4π neutron detector
Search for a flat efficiency 4π neutron detector:
- number of rings
- distance between beam axis and center of each ring
- number of detectors per ring
- size of moderator
GEANT simulations for ~50 possible configurations during 2 months on 36 cores

## Slide 16
Neutron multiplicity sorting with a 4π neutron detector
Flat-efficiency 4π neutron detector -> ε(En) no longer depends on En
Configuration | 3He counters | Distance
Inner ring | 4 | 5.5 cm
Middle ring | 9 | 13 cm
Outer ring | 18 | 16 cm
Total | 31 | 
Moderator size: 46 cm

## Slide 17
Neutron multiplicity sorting with a 4π neutron detector
Detection efficiency:
(,n) neutrons:
ε = 40%
(,2n) neutrons:
Both neutrons detected:
ε2 = 16%
Only one neutron detected:
ε(1- ε) = 24%
(,3n) neutrons:
3 neutrons detected:
ε3 = 6.4%
2 neutron detected:
ε2 (1- ε) = 9.6%
Only one neutron detected:
ε(1- ε)2 = 14.4%

## Slide 18
Neutron multiplicity sorting with a 4π neutron detector
Sensitivity test
Sensitivity to displacements of counters from the inner ring.
Sensitivity to moderator size.

## Slide 19
Neutron multiplicity sorting with a 4π neutron detector
Comparison with calibration source
Calibrated 252Cf source obtained from NIST Japan

## Slide 20: Partial and total photoneutron cross sections for 209Bi
  (g,1n) measurements:
  7.5 to 14.35 MeV -ray beams
  Nd:YVO4 laser INAZUMA
  – 1st harmonic (1064 nm)
  Electron energies - 600 to 900 MeV
  (g,2n), (g,3n), (g,4n) measurements:
  14.5 to 40 MeV -ray beams
  Nd:YVO4 laser INAZUMA
  – 2nd harmonic (532 nm)
  Electron energies - 640 to 1100 MeV
  Flat efficiency 4π neutron detector
  LaBr3 detector for energy monitor
  100% NaI-Tl detector flux monitor
  Monoisotopic thick 209Bi target

## Slide 21
The neutron multiplicity sorting technique relies on the following assumptions, so that we can assume that x neutrons recorded between 2 consecutive gamma bunches are in fact coming from the same reaction:
(1) one gamma ray bunch incident on the target generates no more than one (g,xn) reactions
  perform count rate estimates to find convenient target thickness values depending on the cross section and beam intensity and also beam time estimates for statistics.
(2) the reaction neutrons have enough time between two consecutive gamma bunches to be moderated and recorded by the neutron counters
modify the laser frequency to have at least ~500 us moderation time (2 kHz frequency).
As the laser used at NewSUBARU can not be used at frequencies lower than 15 kHz we had to introduce into the optical setup two additional optical elements:
- a Pockels cell (a crystal that changes the polarization angle of the wave that passes it as a function of the high voltage applied on it)
- and a polarizer.
The polarizer and the Pockels cell were set as to allow passage of laser bunches only if specific high voltage was set on the Pockels cell. We operated the laser using an external trigger signal with 60 us period (~16 kHz frequency). The same trigger signal was downscalled to 480 us (~2 kHz frequency) and sent to the Pockels cell.

## Slide 23: INAZUMA Laser 1064 nm first harmonic – 40 W maximum power 532 nm second harmonic – 15 W maximum power

## Slide 24: Pockels Cell and Polarizer

## Slide 25
Laser power

## Slide 26
TAC spectrum for the first neutron

## Slide 27
TAC spectrum for the second neutron

## Slide 28
TAC spectrum for the third neutron

## Slide 29: Collaboration
D. Filipescu 1,2, H. Utsunomiya3, Vladimir Varlamov4, F. Camera5,6, I. Gheorghe 1,3,7, S. Belyshev4, S. Katayama3, T. Shima8, K. Takahisa8,S. Miyamoto9, S. Amano9, T. Glodariu2, M. Sin7
1Extreme Light Infrastructure – Nuclear Physics (ELI-NP) / Horia Hulubei National Institute for R&D in Physics and Nuclear Engineering (IFIN-HH), 30 Reactorului St., Bucharest-Magurele, 077125, Romania.
2Horia Hulubei National Institute for R&D in Physics and Nuclear Engineering, Bucharest-Magurele, 077125, Romania
3Department of Physics, Konan University, Okamoto 8-9-1, Higashinada, Kobe 658-8501, Japan
4Lomonosov Moscow State University, Department of Physics, Moscow, 119991 Russia
5Dipartimento di Fisica, Università degli Studi di Milano, via Celoria 16, I-20133 Milano, Italy
6INFN, Sezione di Milano, via Celoria 16, I-20133 Milano, Italy
7Faculty of Physics, University of Bucharest, 077125, Bucharest, Romania
8Research Center for Nuclear Physics, Osaka University, Suita, Osaka 567-0047, Japan
9Laboratory of Advanced Science and Technology for Industry, University of Hyogo, 3-1-2 Kouto, Kamigori, Ako-gun, Hyogo 678-1205, Japan

## Slide 30: Experimental team

## Slide 31
ELI-NP Team, March 5, 2013
Thank you!

## Slide 32
Large NaI gamma flux monitor detector pile-up spectrum

## Slide 33
Counters Enery-Time response for 1 MeV neutrons as predicted by GEANT4

## Slide 34
All neutron counters are ORed

## Slide 35
Electronic signals treatment for the first neutron

## Slide 36
Electronic signals treatment for the second neutron

## Slide 37: Calculations of 209Bi(g,Sn) cross section using the Empire code
Gamma strength function: Modified Lorentzian 1 (MLO1) – taken from RIPL GDR parameters - taken from RIPL
- first choice - deduced from experimental data
- second choice - parametrizations.
Default Level density - Enhanced Generalized Superfluid Model (EGSM) from RIPL.
Default optical potential for neutrons and protons is Koning-Delaroche regional potential.
Preequilibrium model - the exciton model implemented in PCROSS.

## Slide 38: Calculations of 209Bi(g,xn) cross section using the EMPIRE code