## Slide 1: Data supply of accelerator devices Data management of device process data at a medical accelerator from an operator´s view
M. Galonska, S. Scheloske, R. Cee, E. Feldmeier, K. Höppner, J. Mosthaf, C. Schömers, A. Peters, T. A. Haberer
and the
HIT team

## Slide 2: The HIT facility
2
Heidelberg Ion Beam Therapy Center:
Dedicated particle/heavy ion therapy facility
3D intensity controlled raster scanning
Scanning heavy ion gantry
Patient treatment since 2009
Gantry clinical operation 2012
R & D in a broad range of disciplines
Notes: 2

## Slide 3: Outline
3
Operator´s tools for tracking, overviewing, and documenting changes of device settings after machine tuning and commissioning (interfaces within and beyond the proprietary control system)
Part of workflow from beam settings (adjustment data) via pre-clinical testing to clinical use of ion beams (therapy data)
Motivated by rapid rise of device data due to commissioning of gantry, oxygen, and helium beams in recent years
Notes: 3

## Slide 4
4
Server architecture, automatic backup and recovery management, data archiving, database operations, disaster recovery, ...
Yes, there is such things at HIT!!!!
Therapy settings (they´re protected against any unintended change, ...at any time! Changes are always instantanously noticed within the control system!)
It´s not about....
Notes: 4

## Slide 5: Pencil Beam Library
5
Pencil Beam Library
ion species | C, p, O, He,... | Steps
energy | 50 – 430 MeV / u | 255
focus | 4 – 13 mm | 4
intensity | 107 – 1010 / spill | 10
gantry angle | 0° − 360° | 36
Accelerator Model
Settings for each combination source-treatment room / ion type provided by control system, i.e. E, F, I, G
Any E, F, I can be requested cycle-by-cycle (treatment plan)
Device settings stored in the device controller units, backuped in a DB
  Therapy, verified flash data: non-volatile memory of controllers (Flash)
  Machine adjustment / experimental data: volatile part of controller (RAM)
Notes: 5

## Slide 6: Data Supply
6
Physical input parameters (accele-
rator model) set to control values
(data supply model)
Settings stored in DB and device
controller for data recall (beam
request)
MIRKO: B. J. Franczak in: Computing in accelerator design and operation, Springer, Berlin, 1984
GUI_Modi: physical input data
Ion optical program MIRKO
Notes: 6

## Slide 7: Data Supply
7
Physical input parameters (accele-
rator model) set to control values
(data supply model)
Settings stored in DB and device
controller for data recall (beam
request)
Proprietary control system supplied by automation company
Oracle database for storing:
  Cycle data
  Measured values (power supplies, beam instrumentation, vacuum pressure, ...)
  Device settings
  Device parameters
  Alarm & error logs
  Devices of virtual accelerators
  Users & access rights
  Computers & applications
Device controller, Eckelmann AG, Wiesbaden, Germany
Notes: 7

## Slide 8
Example MEBT
p
4He2+
12C6+
16O8+
7 MeV / u
Constant Br (LEBT/Linac/MEBT)
One input / ion / source  one set value / ion / source
Export file of physical input data of a
MEBT quadrupole

## Slide 9
9
E x G
E x F x G
Energy and angle dependent settings of
focusing quadrupoles (beam width)
Energy and angle dependent settings of
dipoles/steerer (beam position)
Spline interpolation over base points
energy step
kL [1/m]
Example Gantry
E x G
E x G
E x G
E x F x G-Spline (HEBT): 255 x 4 x 36 input values / ion  255 x 4 x 36 set values / ion
B. J. Franczak, EPAC 96, Spain
Gantry beam line
Notes: 9

## Slide 10: Input Parameters
10
Constant Br (LEBT/Linac/MEBT): one input  one set value / source
Different Br (HEBT): one input  255 set values (implicite E-Dependency), Br-Scaling
Different Br (HEBT), diff. beam target: one input  255 x 5 (ExD-Dependency)
E-Spline (Sync/HEBT): 255 input  255 set values (E-Dependency, some with ExD-Dependency)
ExF-Spline (HEBT): 255 x 4 input  255 x 4 set values
ExG-Spline (HEBT): 255 x 36 input  255 x 36 set values
ExFxG-Spline (HEBT): 255 x 4 x 36 input  255 x 4 x 36 set values
...
 ~500 components, ~180 device controlers
 Times number of ion types!
see for instance: T. Fleck et al., PCaPAC08, Ljubliana, Slovenia
D. Ondreka, EPAC08
Notes: 10

## Slide 11: From settings to therapy
11
  Small facility, but ion beams of different E-F-I have to be provided to all treatment rooms at any time
  Huge amount of device settings, and device data have to be managed
  Manifold options to change device settings (adjustment)
  Maximum flexibility and maximum safety
  Device settings, data have to be carefully tracked and changes have to be analysed before copying adjustment data to therapy data tables...
  ...after beam approval (medical physicist)!
Notes: Ströme E-Septa
11

## Slide 12: Reports
12
Conversation / meetings / discussion
E-Mail
Data collection on a network storage
Written records
Records posted to the Electronic logbook
Special formatted tracking files
.... others are possible (i.e. web based Bugtracker,...)
It´s nonstandard
contains many (all) information, but somehow unsorted
The usual way of communication...
Notes: 12

## Slide 13: Modi Change Log and Data Exports
13
Part of proprietary control system
Log of every single parameter change
Details of change: device, value, time, user, ...
Unfeasible for general view
Historical view ~1 month
Part of proprietary control system
Manual export of input parameters
  MEFI (energy-focus-intensity)
  Virtual Accelerators (beam line)
  Splines
  Control value dumps (output therapy/adjustment data, autom.)
  basedata (calibration data, ...)
Notes: 13

## Slide 14: Data Exports Examples
14
Part of proprietary control system
Manual export of input parameters
  MEFI (energy-focus-intensity)
  Virtual Accelerators (beam line)
  Splines
  Control value dumps (output therapy/adjustment data, autom.)
  Basedata (calibration data, ...)
simple (operator´s) way of backing up input data
simple way of restoring data via file import
not simple way of detect changes / getting general view on modifications due to large number of files
 More than 1,000 files!
Notes: 14

## Slide 15: Data Export Analysis
15
DB scripting (control system experts only)
Stand alone solutions (operator)
  Bash
  File/Directory Compare Software
  It´s nonstandard, but yet very effective
  ....
...but
Analysis tools required
Not implemented in proprietary control system
Joachim Eibl, kdiff3.sourceforge.net
http://www.mingw.org/
Bash MinGW
File/Directory compare software
Notes: 15

## Slide 16: Subversion Repository
16
Subversion repository / daily committed
Not part of proprietary control system
Revision management of all device settings
Ease of use for operator
Accessible to all operators on any control room computer
Standard solution
Clearly arranged display of version changes
  All device settings
  Single device settings
Overview device input data (all devices)
Overview revisions of device input data (all devices)
Implemented by K. Höppner
Notes: 16

## Slide 17
17
Subversion Repository
Revisions (single device)
Old / new data (single device)
Notes: 17

## Slide 18
18
Subversion Repository
Old / new data (single device)
Notes: 18

## Slide 19: Flash Dump Comparator
19
Not part of proprietary control system
Comparison of files (specific formatted files by control system, ~500 MB / file) containing
  Control values
  Firmware/Software versions
  Device parameters / basedata
  IP
  ...all other therapy relevant data!
Generates a result file with clearly arranged display of device / parameter changes
Different level of output (device names and checksums, ..., details of control values)
Indicates potential effect of data change on the beam as decision support for QA procedures
R. Panse, C. Lampe, FDC specifications, and manual
A. Höss et al., Radiation Oncology 2014, 9:79
FDC result file: devices with modified data
FDC result file: details of data changes
 Medical device software used by medical physics staff
Notes: 19

## Slide 20: Flash Dump Comparator
20
Flash Dump Comparator version for operators in control room
Not only used for therapy data comparison, but also for comparing...
  Adjustment data versions
  Therapy and adjustment data versions
R. Panse, C. Lampe, FDC specifications, and manual
A. Höss et al., Radiation Oncology 2014, 9:79
FDC result file: check list as decision support
for QA prodedures
 It´s more than just another file comparator, it´s sophisticated software developed according to medical device standards!!
Notes: 20

## Slide 21: Therapy Data Changes
21
Part of proprietary control system
Changes only in special protected submode
Accelerator and therapy control system
Changes intstantanously visible (accelerator/therapy control system)
Beam request blocked
Analysis by medical physicist
  QA programm
  Approval / denial of approval
K. Höppner, talk coming soon...
Device list with modified therapy data
Shift report:
Device list with modified therapy data
Notes: 21

## Slide 22
22
Thanks!
http://www.klinikum.uni-heidelberg.de
Notes: 22