## Slide 1: Simulations and Studies of ISIS IPMs
Chris Wilcox, Rob Williamson
23/05/2017
christopher.wilcox@stfc.ac.uk, rob.williamson@stfc.ac.uk

## Slide 2
  Layout
  70 MeV H- Linac
  70-800 MeV proton synchrotron
  Two extracted proton beam transfer lines (EPBs) to targets
  Highlighted Parameters
  50 Hz repetition rate
  Intensity ~2.5x1013 protons per pulse
  Vacuum pressure ~10-7 mbar
  At extraction:
    2x100ns proton bunches, ~200ns separation
ISIS Spallation Neutron & Muon Source

## Slide 3: IPMs at ISIS: Transverse Plane
  Profile measurements use positive residual gas ions, not electrons.
  Ions are detected by an array of Channeltrons (electron multipliers).
  No shaping field electrodes.
BEAM
Transverse cross-section view of an ISIS IPM
Drift field electrode (variable up to 60 kV)
40 Detector Array
240 mm
+15 kV (typical)
220 mm

## Slide 4
  Electron multipliers, manufactured by Photonis.
  240mm wide array of 40 Channeltrons, providing 6 mm resolution.
  40 power supplies used to bias each Channeltron individually.
A single channeltron
IPMs at ISIS: Multi-Channeltron Array
Raw beam profile data (taken when all channeltrons have same bias voltage)
The 40 channeltron array used to measure beam profile
6 mm
  With equal biases, gain varies significantly between Channeltrons.
    In-situ calibration required

## Slide 5: IPMs at ISIS: Longitudinal Plane
Simplified longitudinal view of an ISIS IPM
Compensating Field Electrode
Single Motorised Channeltron
(SCPM)
Multiple Channeltron Array (MCPM)
Drift Field Electrode
y
z
  Each IPM contains two sets of detectors and two separate electrodes:
    Compensating field ensures no net kick on the beam
    Single, moveable Channeltron used to calibrate the Multi-Channeltron array.
Raw and Calibrated Beam Profile Data
BEAM
BEAM

## Slide 6
Measurement Errors
  Three main sources of measurement error:
  Transverse drift field shape
  Beam space charge field
  Longitudinal saddle point
Transverse electric potential distribution created by the drift field electrode
Widens the measured profile.
Introduces roughly 30% error in normal operation (15 kV).
Detector Array

## Slide 7
Measurement Errors
  Three main sources of measurement error:
  Transverse drift field shape
  Beam space charge field
  Longitudinal saddle point
Also widens the measured profile.
Effect varies with beam properties and drift field strength.
Time-varying: only present while each proton bunch is passing through the IPM.
Transverse space charge field pattern created by the beam
Detector Array

## Slide 8
Measurement Errors
  Three main sources of measurement error:
  Transverse drift field shape
  Beam space charge field
  Longitudinal saddle point
Longitudinal electric potential distribution created by the two drift field electrodes
Saddle point
MCPM Detectors
SCPM Detector
Caused by the compensating field interacting with the drift field.
Extra ions reach the detectors, following complex trajectories.
The monitor is not sampling a purely 2D transverse slice of the beam.

## Slide 9
Simulation output showing sample 3D trajectories of ions that reach the detectors in the longitudinal plane.
Measurement Errors
  Three main sources of measurement error:
  Transverse drift field shape
  Beam space charge field
  Longitudinal saddle point
Caused by the compensating field interacting with the drift field.
Extra ions reach the detectors, following complex trajectories.
The monitor is not sampling a purely 2D transverse slice of the beam.
MCPM Detectors
Combined effect of all error sources means true beam width can be only ~35-60% of the measured beam width!

## Slide 10
Benchmark Measurements:
EPB1 Monitor
  Software correction is used to compensate for drift field and space charge errors.
    Works reliably for low intensity, centred beams.
  For further study, we moved an IPM to a beam transfer line (EPB1), to take test measurements where the beam is well defined.
SEM Grid Profile Monitor
SEM Grid Profile Monitor
Dual-Plane Position Monitor
Benchmarking IPM
Test IPM location in EPB1. Nearby monitors provide accurate beam parameters for comparison with IPM measurements.

## Slide 11
IPM Simulation Approach
  A thorough IPM simulation model is in development, and is benchmarked against measurements from the EPB monitor.
    Aim to verify the profile correction scheme and expand it to varied beam properties.
3D Electric Field Calculation (CST EM Studio)
Calculation of drift fields and beam field
Two fields calculated – with and without beam
Live Beam Measurements (ISIS)
When benchmarking
Detector array of 40 channeltrons
Compensating field hv plate
Single channeltron detector
Drift field hv plate
Beam
Simplified monitor geometry used in the CST model

## Slide 12
IPM Simulation Approach
  A thorough IPM simulation model is in development, and is benchmarked against measurements from the EPB monitor.
    Aim to verify the profile correction scheme and expand it to varied beam properties.
3D Electric Field Calculation (CST EM Studio)
Calculation of drift fields and beam field
Two fields calculated – with and without beam
Particle Tracking Simulation (C++)
Creates and tracks ion distribution through E-field
Time-varying E-field approximation
Post Processing (IDL)
Calculates profile measurement
Applies weightings to account for detector behaviour
Data Analysis & Visualisation (IDL)
Detector Characterisation Experiments
Vacuum tank & ion gun
Live Beam Measurements (ISIS)
When benchmarking
Acceleration calculated from Lorentz force on each ion, B≡0.
Tracking using 2nd Order Euler Method, 1 ns time step
*see appendix for detail

## Slide 13
Simulated Profiles Vs. Measurement
  Simulation output profiles closely match measurements from the test IPM.
    …when running at low beam intensities.
Comparisons between real IPM beam profile measurements (black) and simulation results (blue) for the low beam intensity measurements (~5x1012 ppp)

## Slide 14
  At high intensities, simulation does not reproduce high enough amounts of profile widening.
  Further studies to understand why:
    Channeltron detector studies in lab test setup.
    Machine measurements with a fast amplifier, to study measured profile time structure.
Simulated Profiles Vs. Measurement
  Simulation output profiles closely match measurements from the test IPM.
    …when running at low beam intensities.
A typical high intensity beam profile (2.5x1013 ppp) comparison between simulation and measurement.

## Slide 15
Sample trajectories of ions that reach the detectors in the longitudinal plane.
Channeltron Tests: Motivation
  Simulation output showed ions entering the Channeltrons at different angles of incidence.
    Detection efficiency vs. incident angle data was not available.

## Slide 16
Vacuum Tank Tests: Setup
Ion gun
Electron gun
Phosphor screen
Channeltrons
Motorised mounting plate
Channeltron test housing inside the vacuum tank. Test housing geometry matches the IPM Channeltron housing.
Faraday cup

## Slide 17
Channeltron Tests: Motivation
  Simulation output showed ions entering the Channeltrons at different angles of incidence.
    Detection efficiency vs. incident angle data was not available.
    Neither was accurate data for variations in ion energy or mass.
  Photonis advised us to test the Channeltrons ourselves, and not rely on “approximate” data available in their brochure.
Histogram showing the angles of incidence at which the simulated ions enter with the detectors.
Longitudinal Angle of Incidence
No. ions Detected
Saddle point ions ~30-50o
“Normal” ions
~5-20o

## Slide 18
Vacuum Tank Tests: Results
  Channeltrons bombarded at various angles with a 4 keV Helium ion beam.
    Higher angle saddle point ions detected at ~50% of peak efficiency.
    Weightings calculated from these results and applied in simulation.
Measured detection efficiencies of multiple array Channeltrons vs. ion angle of incidence

## Slide 19
Vacuum Tank Tests: Results
  Channeltrons bombarded at various angles with a 4 keV Helium ion beam.
    Higher angle saddle point ions detected at ~50% of peak efficiency.
    Weightings calculated from these results and applied in simulation.
  Further tests:
    Different models of Channeltron produce different results (SCPM vs. MCPM).
    Channeltron behaviour more stable with higher bias voltages.
    Detection efficiency decreases as ion mass increases.
  Advice: make sure you understand your detectors when simulating an IPM!
Comparison of detection efficiencies of the two different models of detector used in ISIS IPMs

## Slide 20
Fast Amplifier Measurements
  Recent studies taking fast amplifier measurements from our benchmarking IPM.
    Profile time structure provides a different comparison between measurement and simulation.
  Compensating electrode switched off to remove saddle point and simplify behaviour.
Comparison of profile time structure when measured with standard amplifier (left) and with a fast amplifier (right)
Time (ns)
0
5000
Current Measured on each Channeltron
Measurement with Standard 2.5 Ms/s Amplifier
Time (ns)
0
4000
Current Measured on each Channeltron
Measurement with Fast 60Ms/s Amplifier

## Slide 21
Typical time structure from the benchmark IPM containing 5 peaks, seen when a fast amplifier is used for measurement
Time
Channeltron Current
Fast Amplifier Measurements:
Multiple Time Peaks
  Fast amplifier measurements detect five separate peaks of signal, at different times.
    Saturation of the 3rd or 4th peak (depending on bias settings) probably the reason for mismatch between simulated and measured profiles at higher beam intensities.
Typical difference between simulated and measured profiles at high beam intensities

## Slide 22
Typical time structure from the benchmark IPM containing 5 peaks, seen when a fast amplifier is used for measurement
Time
Channeltron Current
Fast Amplifier Measurements:
Multiple Time Peaks
  Time structure in simulation matched the shape and timings of the first two peaks.
    Initial thought: 1st and 2nd peak contain true profile data. Later peaks are noise from the downstream neutron target.
  However, test measurement where one pulse deliberately missed the target still showed the same 5 peaks.
    All 5 peaks must be related to the beam.
Simulated time structure of a single bunch. In this result the residual gas was modelled as 100% hydrogen.
~150 ns

## Slide 23
  Individually, each of the 5 peaks produce similar profile data.
    Additionally, the peak times increase if drift field strength is decreased.
  All 5 peaks likely to be caused by residual gas ions created by the beam.
    Each peak could contain valid profile information, can we prove what causes them?
    If they each contain true profiles, we can set the IPM up to use specific peaks only for measurement.
Chart: Integrated Channeltron Signals vs. Time: Low Channeltron Bias & Low Beam Intensity
With a lower Channeltron bias, the strong peaks do not saturate but the 1st and 2nd peaks are too weak to take a measurement from.
Fast Amplifier Measurements:
Multiple Time Peaks

## Slide 24
Simulation:
Multiple Residual Ion Species
  We now suspect the 5 peaks are caused by different residual gas ion species.
  Old mass spectrometer measurement of beam vacuum suggests a strong O and CO (or N2) presence.
    Adding these different ions to the simulation generates the 5-peak time structure (approximately), but the different peak heights vary significantly.
  This is a work in progress, accurate residual gas distribution not yet known.
    Probabilities of ionisation of each species also unknown.
    Detection efficiency variation of Channeltrons for different ion masses also needs to be measured.
Chart: High bias, low intensity time structure
Comparison of measured time structure (left) and simulated time structure (right). The peak heights are not yet modelled accurately in the simulation.

## Slide 25
Future Work
  Initial tests in vacuum tank showed that an argon ion beam was detected at ~60% efficiency of a helium ion beam.
    Need more complete experiment results.
    What is the probability of ionisation for each residual gas species?
  Further comprehensive Channeltron characterisation experiments planned.
    Detection efficiency against varying Channeltron bias and ion energy
  Future monitor design can take into account what we are learning.

## Slide 26
Future Work
  Can we operate the IPM with the compensating field off permanently?
  With the compensating field on, the time structure has additional features:
  Recent fast amplifier measurements on a synchrotron ring IPM give the same time peak structure, but only at extraction.
    Why? Further studies required.
Compensating Field On
Compensating Field Off
Fast amplifier measurements of beam profile, showing extra detail when the compensating field is switched on.

## Slide 27: Thank you for your time
Questions?
Compensating Field On
Compensating Field Off

## Slide 28: Additional Slides

## Slide 29
C++ Particle Tracker:
Equations Used to Calculate Ion Motion
= 0
Lorentz force acting on each ion
Ion velocity change over 1 time step
v = final velocity, u = initial velocity
Ion position change over 1 time step
x = final position, x0 = initial position
Final time step length when ions reach the detector
The final y-position is known (the detector position), so the ion is tracked over a smaller final time step, tf . This give a precise ion position and velocity as it reaches the detector.
Solve as a quadratic of tf .
Use the +ve solution as the final time step length. Recalculate final velocity and position in each plane using the new time step.

## Slide 30: Effect of beam on IPM E-Fields
Longitudinal potential distribution once the beam has left the monitor
Longitudinal potential distribution with the beam present in the monitor
Transverse potential distribution once the beam has left the monitor
Transverse potential distribution with the beam present in the monitor

## Slide 31: Experimental Fast Amplifier Data
Data showing a sweep across different channeltron bias voltages, causing changes to time structure peak heights.
Chart: Time (ns)

## Slide 32: Profile Correction Scheme
Derived from combination of simulation and measurements
Correct synchrotron profiles to within 6 mm
Profile tails
Ionisation monitor measured width and true beam width at percentage p
Drift field voltage
Drift field correction
Space charge correction
Relativistic gamma
Beam intensity
R. Williamson
EPAC ‘08
Ideal profile (purple)
Simulated profile (red)
Corrected profile (blue)
Simulated data

## Slide 33: Time Structure: Compensating Field Off vs. Compensating Field On
Comp. Field On
Comp. Field On

## Slide 34
Comparison With Other Codes
  From “Ionization Profile Monitor Simulations – Status and Future Plans” M. Sapinski et al., IBIC2016
3D

## Slide 35: Ion Gun Specs
HIDEN Analytical
IG20 Oxygen/Gas Gun

## Slide 36
Profile Simulator Workflow: Setup
Inputs
(Data from ISIS machine run)
3D E-Field
Calculation
(CST EMS)
Ion Tracking Setup
(C++)
CST EM Studio IPM Models – Calculate E-Fields within monitor
2 CST IPM models are used for time dependence approximation:
1 model with beam inside monitor, 1 model without beam
C++ Tracking Code Setup
Import E-Field values from the CST model with beam inside the monitor.
Create uniformly spaced ion distribution, of various masses, within beam volume then randomise initial positions
Initial ion velocities are set to 0.
Export E-Fields as 3D grids with data point spacing of 2mm in each plane
Data point spacing <1% of IPM detector width (240mm).
Begin tracking loop
Beam profile information from 2D SEM grid
Horizontal & vertical radii, positions.
Beam intensity
IPM Drift Field Strength
User Inputs
Beam radii & position
Residual gas ion mass & charge
Tracking region boundaries (detector & monitor wall positions).

## Slide 37
Profile Simulator
Workflow: Ion Tracker
Ion Tracking Loop
(C++)
Loop for each ion in the distribution (i)
Calculate nearest E-Field matrix point and retrieve field values
No interpolation of E-Field used as point spacing <1% of detector width.
Calculate ion motion during time step
New position and velocity calculated using initial position, initial velocity and Lorentz force applied by E-Field.
Loop for each time step (t) until all ions have left tracking region
Time step size, Δt = 1 ns
Has ion
left tracking region*?
No
Yes
i=i+1
Calculate final time step size, Δtf
If the ion reached the detector, calculate exact time of arrival (<1ns). Use this new time step to recalculate exact position & velocity at contact with detector.
Exclude ion from further tracking loops
Have all ions been tracked in current time step?
No
Yes
t=t+Δt
*The user defines the “tracking region” during setup. The tracking region represents the volume inside the monitor, with boundaries on the ion detectors and the monitor walls.
If t=200 ns, load the beam independent E-Field matrix
Replace the initial E-Field matrix, calculated with the beam inside the monitor, with the E-Field from the CST model with no beam present.

## Slide 38
Profile Simulator Workflow:
Post Processing
Simulation Post Processing
Code
(IDL)
Load ion motion tracks calculated in C++ ion tracker code
Create simulated profile measurement from each ion that has a final position on the detector face
Measured profile consists of a histogram with 40 bins, each with a width of 6mm. Each bin represents one channeltron detector.
Apply an elliptical weighting to the measured profile, using the initial (ionisation) positions of each ion
The ion tracker created a uniform distribution of ions within the beam volume, but in reality the ion creation will more concentrated in the beam centre.
Beam distribution weighting
Weighting to simulate detector behaviour
Apply a weighting to simulate the behaviour of the channeltron detection behaviour
The channeltron detection efficiency decreases as the angle of incidence increases. It also varies with the energy of the incoming ion.
Apply space charge and drift field correction
The profile data without this correction applied should match the IPM measured profile.
The profile data with this correction applied should match the SEM grip measured profile.
Calculate & output results