## Slide 1
The KATRIN Experiment
Joachim Wolf
KIT, 09.03.2017
KATRIN Experiment
Tritium Source
Pumping and Transport Section
Spectrometer and Detector Section
Conclusions / Next Steps

## Slide 2
The KArlsruhe TRItium Neutrino Experiment
Goal: measure the effective neutrino mass
International KATRIN collaboration:
  about 130 members
  6 countries (GER, US, CZ, RUS, ES, F)
  19 institutions
Sensitivity on m(νe): 2 eV/c2 → 0.2 eV/c2

## Slide 3
3H: super-allowed | 
E0 | 18.6 keV
t1/2 | 12.3 y
Source & Transport Section (STS)
Spectrometer & Detector Section (SDS)
ideal
b-emitter
most sensitive method
G. Drexlin, V. Hannen, S. Mertens, C. Weinheimer, Current Direct Neutrino Mass Experiments (Review), Adv. In High Energy Phys. (2013) 293986
The KArlsruhe TRItium Neutrino Experiment
70 m
J. Wolf - The KATRIN Experiment

## Slide 4
  collimation: μ = E⊥ / B = const,  E⊥ → E|| for B = 6 T → 3mT
  energy analysis: transmission condition: E|| > eUmax (retarding potential)
  energy resolution: ΔE = E ∙ Bmin / Bmax = 18.6 keV ∙ 0.3 mT / 6 T = 0.93 eV
Magnetic Adiabatic Collimation
with Electrostatic Filter
The MAC-E Filter
A. Picard et al., NIM B 63 (1992)‏

## Slide 5
KATRIN overview: 70 m beamline
windowless gaseous
tritium source
differential
pumping
cryogenic
pumping
integral
energy
analysis
electron
counting
diagnostics

## Slide 6
KATRIN overview: particle tracking
ultra-precise calculation of electro- / magneto-static fields
high-statistics simulations of particle trajectories
Kassiopeia software also used by other projects

## Slide 7
tritium source: 1011 b-decays/s
total background: 10-2 cps
 10-3 stability of tritium source column density
10-3 isotope content in source
 10-5 non-adiabaticity in electron transport
 10-6 monitoring of HV-fluctuations
 10-8 remaining ions after STS
 10-11 mbar in Main Spectrometer
 10-14 remaining flux of molecular tritium in SDS
experimental challenges
Source & Transport Section (STS)
Spectrometer & Detector Section (SDS)
KATRIN overview: benchmark parameters

## Slide 8
KATRIN Source and Transport Section

## Slide 9
KATRIN Source and Transport Section
 Source - stringent control of systematic effects: isotopic purity,
gas temperature, plasma effects

## Slide 10
Beam tube temperature
Technological development
Challenge
temperature stability on 10-3 level
- novel 2-phase neon cooling system
- required: ΔT = ± 30 mK (1 h) - achieved: ΔT = ± 1.5 mK (1 h)
→ stability surpassing specifications
beam
tube Temp.
2-phase
neon
Windowless Gaseous Tritium Source WGTS

## Slide 11
Windowless Gaseous Tritium Source WGTS
- calibrated Laser-Raman system
for all 6 hydrogen isotopologues
- achieved: < 10-3 accuracy in 60 s
Raman spectroscopy
Technological development
Challenge
- measure isotopic source content
with 10-3 accuracy in 100 s
isotopic
content
Raman
spectr.

## Slide 12
tritium free
Chart: 
T2 injection
differential & cryo-pumping
101
10-1
10-3
10-5
10-7
10-9
10-11
10-13
10-15
10-17
tritium flow rate [mbar ℓ /s]
Suppress tritium flow from source to spectrometer by factor > 1014
differential pumping
cryogenic
pumping
Transport and Pumping Sections

## Slide 13
Transport and Pumping Sections
DPS2-F
b
Differential Pumping Section (DPS)
active pumping: 6 + 8 + 4 TMPs
tritium retention: 107
tritium ion removal: dipole electrodes
magnetic field: 5.6 T
built at KIT, commissioning 2016
Cryogenic Pumping Section (CPS)
based on by cryo-sorption
tritium retention: >107
beam diagnostics: retractable detector
magnetic field: 5.6 T
delivered: 2015, commissioning: 2016
b
Argon Frost Pump
T = 3 – 4.5 K
T2
cryo-sorption
Ar Frost
stainless steel
CPS
F. Eichelhardt et al, Fusion Science and Tech. 54 (2008) 615
DPS1-R
DPS1-F
X. Luo et al., Vacuum 80 (2006), 864 - 869

## Slide 14
Risky operating conditions for TMPs?
TMP in a magnetic field
eddy currents can over-heat rotor
high mag. field can slow down rotor
failure of magnetic bearing
test setup built at KIT for large TMPs
math. model developed for prediction
Endurance test for TMP with tritium
tritium can affect non-metal parts of pump
TMP type: Leybold MAG-W 2800
tested at Tritium Laboratory Karlsruhe (TLK)
398 days operation with tritium
throughput: 1106 g tritium
F. Priester, PhD thesis at KIT (2013)
R. Größle et al., Vacuum 86 (2012) 985-989

## Slide 15
Complete dismantling of a MAG W 2800
parts were highly contaminated with tritium, but …
… parts looked like new, no indication of wear, cables and O-rings ok

## Slide 16
Risky operating conditions for TMPs?
TMP in a magnetic field
eddy currents can over-heat rotor
high mag. field can slow down rotor
failure of magnetic bearing
test setup built at KIT for large TMPs
math. model developed for prediction
Endurance test for TMP with tritium
tritium can affect non-metal parts of pump
TMP type: Leybold MAG-W 2800
tested at Tritium Laboratory Karlsruhe (TLK)
398 days operation with tritium
throughput: 1106 g tritium
F. Priester, PhD thesis at KIT (2013)
R. Größle et al., Vacuum 86 (2012) 985-989

## Slide 17: TMP in a magnetic field
Pyrometer
Helmholtz coils
( = 120 cm)
TMP
gas inlet
Helmholtz coils: radius = 60 cm
B-field: 0 – 50 mT
coils can be turned by 90°
pyrometer used for rotor temperature
gas flow possible
measures parameters for empirical model

## Slide 18: TMP in a magnetic field
WGTS: TMPs operated in 18.5 mT
8 mm St37 steel shielding: < 70°C
Bmax at 70°C rotor temp.
no shielding: 2 mT
5 mm shield.: 13 mT
8 mm shield.: 20 mT

## Slide 19
The KATRIN Main Spectrometer

## Slide 20
  MAC-E Filter principle → precise electron energy measurement
    Vacuum vessel & electrodes on
    variable retarding potential (18.6 kV)
    Magnetic guiding field: 0.3 mT – 6 T
    High resolution: ΔE = 0.93 eV @ 18.6 keV
KATRIN Main Spectrometer
windowless gaseous electron transport and
tritium source pumping section
Main Spectrometer
detector
arrival at KIT: 26.11.2006

## Slide 21
KATRIN Main Spectrometer
arrival at KIT: 26.11.2006
  MAC-E Filter principle → precise electron energy measurement
    Vacuum vessel & electrodes on
    variable retarding potential (18.6 kV)
    Magnetic guiding field: 0.3 mT – 6 T
    High resolution: ΔE = 0.93 eV @ 18.6 keV
  Stainless steel (~200 to, 316LN)
  Dimensions:
    diameter: 10 m
    Length: 23 m
      volume: 1240 m³
      inner surface: 1240 m² (including wire electrodes)
windowless gaseous electron transport and
tritium source pumping section
Main Spectrometer
detector

## Slide 22
KATRIN Main Spectrometer
arrival at KIT: 26.11.2006
  Requirements:
  Low pressure (~ 10-11 mbar)
    tritium partial pressure <10-21 mbar
    few radon decays per day
    outgassing rate ~10-12 mbar·ℓ/s·cm2
    total leak rate < 5·10-9 mbar·ℓ/s
  Bakable at 350°C (NEG activation)
  Stable operation at 20°C
windowless gaseous electron transport and
tritium source pumping section
Main Spectrometer
detector

## Slide 23
KATRIN Main Spectrometer
arrival at KIT: 26.11.2006
  Requirements:
  Low pressure (~ 10-11 mbar)
    tritium partial pressure <10-21 mbar
    few radon decays per day
    outgassing rate ~10-12 mbar·ℓ/s·cm2
    total leak rate < 5·10-9 mbar·ℓ/s
  Bakable at 200°C (NEG elec. activated)
  Stable operation at 20°C (10°C)
  Vacuum components operated in
    Magnetic field: 0.3 mT – 6 T
    Electric potential: 18.6 kV
windowless gaseous electron transport and
tritium source pumping section
Main Spectrometer
detector

## Slide 24
Roughing pump: 640 m3/h screw-pump
6 turbo-molecular pumps (Leybold MAG-W 2800): 10 000 ℓ/s (H2)
KATRIN Main Spectrometer Vacuum
3 TMPs
3 TMPs

## Slide 25
Roughing pump: 640 m3/h screw-pump
6 turbo-molecular pumps (Leybold MAG-W 2800): 10 000 ℓ/s (H2)
Fore-vacuum: 300 ℓ/s TMP and scroll pump (30 m3/h)
KATRIN Main Spectrometer Vacuum

## Slide 26
Roughing pump: 640 m3/h screw-pump
6 turbo-molecular pumps (Leybold MAG-W 2800): 10 000 ℓ/s (H2)
Fore-vacuum: 300 ℓ/s TMP and scroll pump (30 m3/h)
3 NEG-pumps (3000 m SAES St707 getter strips): ~106 ℓ/s (H2)
KATRIN Main Spectrometer Vacuum

## Slide 27
Roughing pump: 640 m3/h screw-pump
6 turbo-molecular pumps (Leybold MAG-W 2800): 10 000 ℓ/s (H2)
Fore-vacuum: 300 ℓ/s TMP and scroll pump (30 m3/h)
3 NEG-pumps (3000 m SAES St707 getter strips): ~106 ℓ/s (H2)
3 cryogenic LN2 baffles (radon): ~160 000 ℓ/s (Rn)
KATRIN Main Spectrometer Vacuum
250 000 ℓ/s
2

## Slide 28
248 wire electrodes on the inner surface
  23 440 insulated wires
  120 000 individual parts
Installed under cleanroom conditions
2008 – 2012: Wire Electrode Installation

## Slide 29
2012: Preparation of first measurements
Electrode installation completed
Vacuum chamber closed
Successful leak test
first bake-out, NEG activ. in Jan. 2013
> 1 year of data from 2013 - 2016

## Slide 30
Indicator for NEG activation:
  pressure ratio between vessel and pump port
  after baking: pP3/pF9 = 0.41  SNEG ≈ 290 m3/s
NEG activation
F9
P3
MolFlow+ simulation
Expected outgassing rate: 10-12 mbar·ℓ/s·cm2
Actual H2 outgassing rate: ~1.4 – 2.5·10-12 mbar·ℓ/s·cm2
Surface (vessel & electrodes): 1240 m2
Effective pumping speed: ~290 m3/s
Expected pressure in vessel: 3.3·10-11 mbar
Actual pressure in vessel: ~1.3·10-10 mbar
M. Arenz et al., JINST 11 (2017) P04011

## Slide 31
 main spectrometer works as high-resolution MAC-E filter:
- sharp transmission function for 18.6 keV e- from egun
- width limited by egun emission spectrum
- HV stability on ppm-scale
e-gun with small angular spread
transmitted e-
reflected e-
KATRIN - MAC-E filter characteristics
different
emission
angles

## Slide 32
Radon as background source (problem)
F.M. Fränkle et al., Astropart. Phys. 35 (2011) 128
S. Mertens et al., Astropart. Phys. 41 (2013) 52
NEG
219Rn emanation from St707 NEG getter strips (2000 m) in pump ports
220Rn emanation from stainless steel walls/weldings
Auger
electrons
shake-off
electrons
conversion
electrons
shell reorganisation electrons

## Slide 33
 passive background reduction: LN2-cooled baffles to cryo-sorb 219Rn, 220Rn
Radon as background source (solution)
Reduction of effective NEG pumping speed: 38%
Reduction of 219Rn flow into main vol. : ~ 0.6%
Reduction of 220Rn flow into main vol. : ~ 6%
G. Drexlin et al., Vacuum 138 (2017), 165–172
MolFlow+ simulation

## Slide 34
 passive background reduction: LN2-cooled baffles to cryo-sorb 219Rn, 220Rn
Radon as background source (solution)
Reduction of effective NEG pumping speed: 38%
Reduction of 219Rn flow into main vol. : ~ 0.6%
Reduction of 220Rn flow into main vol. : ~ 6%
Pumping speed for 219Rn from walls: 160 000 ℓ/s
Pumping speed for 220Rn from walls: 75 000 ℓ/s
G. Drexlin et al., Vacuum 138 (2017), 165–172
MolFlow+ simulation

## Slide 35
Remaining background (0.5 cps) ?
 H* Rydberg atoms:
- desorbed from walls due to 206Pb recoil ions
- non-trapped electrons on meV-scale
- bg-rate: ~0.5 cps
- countermeasures (work in progress):
reduce H-atom surface coverage:
- extended bake-out phase
- strong UV illumination source
isotropic bg for
longer exposure
< 100 meV
H*
210Pb decay

## Slide 36: 238U decay chain (222Rn, 210Pb, 206Pb)
www.wikipedia.com (Tosaka)

## Slide 37: KATRIN milestone 2016 – first light

## Slide 38: KATRIN – next steps
final commissioning tests 2017
spectrometer bake-out 2017
more engineering runs with deuterium 2017/18
first tritium data 2018
operation: ~ 5 years

## Slide 39
Thank you for your attention
Supported by:

## Slide 40: Backup slides

## Slide 41
The KATRIN Setup - Overview
Tritium source
Transport section
1010 e- /s
e-
3He
Pre spectrometer
103 e- /s
e-
e-
e-
Spectrometer
1 e- /s
e-
Detector
ΔE = 0.93 eV
3H
E > 18.3 keV
≈ 70 m
E = 18.6 keV
1011 e- /s
e-
3H
β decay
3He

## Slide 42
Magnetic Adiabatic Collimation
with Electrostatic Filter
The MAC-E Filter
A. Picard et al., NIM B 63 (1992)‏
isotropically emitted
tritium b-electrons
adiabatically collimated
by magnetic field
electrons filtered
by electric potential
remaining electrons
counted after filtering
eU0

## Slide 43
column
density
electron
gun
activity
X-ray
detector
isotopic
content
Raman
spectr.
gas
injection
pressure
stabil.
beam
tube T
2-phase
neon
beam
size
Si PIN
det.
ion
diagnost.
FT-ICR
Source systematics
Challenge Technological development
reduce source systematics monitoring & control of all relevant by > factor 10 compared to parameters with unprecedented previous experiments precision & accuracy
Source and Transport Section
ion
removal
dipole
electr.
CMS DPS1-R WGTS DPS1-F DPS2-F CPS
calibration, differential windowless gaseous tritium differential pumping cryogenic
monitoring pumping source pumping

## Slide 44
DPS 2-F – differential pumping section
subsections tilted by 20°
TMP
CPS
WGTS
 DPS: active differential pumping by 4 main TMPs - retention factor 105

## Slide 45
CPS
 CPS: passive cryotrap
CPS – cryogenic pumping section
adiabatic guiding of electrons
- 7 s.c. solenoids (B = 5.6 T)
- 3 K beam tubes with Ar frost
- tritium retention factor > 107
cryogenic pumping
Port instrumentation
vertical access port for
condensed 83mKr source
horizontal port for monitoring

## Slide 46: Model 1: fit of parameters k1 … k6
pfv
Leybold MAG W 2800
TR (measured)
TR (fit)
B

## Slide 47: Model 1 test: Rotor temperature in a pulsed mag. field
Leybold MAG W 2800

## Slide 48
Influence of magnetic field on TMP
B0 · sin(φ)
φ
parallel field:
failure of magn. bearing (PZ12)
  for B ↑ at 12.6 mT
  for B ↓ at 21.5 mT
no heating of the rotor
perpendicular field:
heating of the rotor (eddy current) < 5 mT
reduction of rotation speed at 8 - 10 mT
bearing stable up to 40 mT
controller in magnetic field:
fan fails at 6.5 mT
pump shut down at 11 mT
 TMPs need magnetic shielding at WGTS and DPS
(magnetic field values valid for MAG W 2200 and 2800)

## Slide 49
potential [kV]
-20
-10
0
1
10-1
10-2
1:20000
Bmin
Bmax
B-field [T]
F. Glück, Prog. in Electromagnetics Research B, 32 (2011) 351-388 & 319-350
70 m
Magnetic field & electrostatic potential
tritium source
spectrometer

## Slide 50
2013: Spectrometer Commissioning
Vacuum conditioning for the MAC-E-filter test measurements
  Plan: baking of the M.S. at 350°C (cleaning and activation of NEG)
  Goal: reach UHV conditions with p ≈ 10-11 mbar
  Bake-out in January 2013
Problems during bake-out (partly solved)
  Short circuit between current leads to electrodes @ 200°C – 300°C
    Reduced baking temperature (300°C) to avoid further damage
  Leakage in CF flange at 50°C
    Differentially pumped vacuum sleeve
    Another leakage after Ar venting
  Leakage in beam-line valve
    Ar venting for repair
Detector and e-gun connected
Start of MAC-E-filter tests
temperature map

## Slide 51
KATRIN Main Spectrometer and Detector Commissioning 2013
angular selective electron gun
transmission function properties
UHV
conditions
10-11 mbar
6 weeks baking,
Tmax = 300 °C
Si-PIN
detector
low background,
148 pixel
MAC-E filter,
energy analysis
vessel on HV,
wire electrode

## Slide 52
KATRIN Main Spectrometer (Deggendorf)

## Slide 53
340 km
8800 km
↕ 7 cm
KATRIN Main Spectrometer Journey to KIT

## Slide 54: Baking of the Main Spectrometer
Duration of baking: 4 weeks
24/7 shifts
Max. temperature: 300°C (now 200°C)
Heating rate: 1°C – 5°C/h
Thermal expansion during bake-out: ~ 10 cm

## Slide 55
Spectrometer Commissioning: Bake-out

## Slide 56
Vacuum status after bake-out

## Slide 57
Coupling of Spectrometer and Detector
    Detector de-coupled during bake-out
    Requires valve inside magnet bore
    O-ring partly slipped out during baking
    Challenge: attach detector without saturation of the activated NEG-pump
X-rays
visual

## Slide 58
1240 m³
Ar 6.0
XENON 1t
gas purification system (SAES)
144 bottles Argon N6.0
O-ring exchanged in Ar atmosphere
beam-line valve now leak tight
detector section attached
    Solution: replacing the O-ring under inert gas atmosphere (Ar)
    Gas quality N9.0 required to prevent contamination of NEG
Ar 9.0
Coupling of Spectrometer and Detector

## Slide 59
Vacuum status after venting with argon

## Slide 60
Indicator for NEG activation:
  pressure ratio between vessel and pump port
NEG activation
F9
P3
simulation

## Slide 61
KATRIN Main Spectrometer Measurements
mean global rate: 0.78 ± 0.20 cps
mean global rate: 0.47 ± 0.09 cps
detailed transmission and background studies
sharpest transmission function ever measured with MAC-E filter
background from 219Rn/220Rn emanation eliminated
will be improved during 2014 commissioning runs

## Slide 62: Vacuum scheme of the Main Spectrometer

## Slide 63: Flanges and Gaskets:
UHV:
  CF gaskets up to 250 mm
  HTMS double gaskets:
    500 mm ground-electrodes
    1700 mm pump ports
  all gaskets bakable at 350°C
intermediate vacuum:
  CF gaskets
fore-vacuum:
  KF O-rings
  ISO K for pump-down and venting
KATRIN Main Spectrometer Vacuum

## Slide 64: KATRIN Main Spectrometer
Component | Material | Temp. | Surface
Main Spectrometer vacuum vessel | 316LN | 20°C | 690.0 m2
Wires (23440 wires with a total length of 42400 m) | 316L | 20°C | 33.6 m2
Electrode frames (248 modules) | 316L | 20°C | 436.8 m2
Electrode rail system | 316LN | 20°C | 58.0 m2
Feedtrough flanges | 316LN | 20°C | 2.0 m2
Small components (frame NEG-pumps, etc.) | 316L | 20°C | 1.5 m2
Σ stainless steel | 316L(N) | 20°C | 1221.9 m2
Σ ceramic insulators | Al2O3 | 20°C | 5.8 m2
Σ anti-penning electrodes | Ti | 20°C | 11.0 m2
Σ ground electrodes | Al | 20°C | 1.3 m2
Σ surfaces at room temperature |  | 20°C | 1240 m2
Σ cryogenic baffles | Cu | 77 K | 31 m2
Σ NEG-strips | St707 | 20°C | 180 m2
Volume Main Spectrometer |  |  | 1240 m3

## Slide 65: Simulations of the Main Spectrometer
simplified model of the main spectrometer created (optimized discretization for Molflow)
simulate pressure ratio pP3 / pF9 of pressure gauges

## Slide 66
three possible gas sources for hydrogen and radon:
  complete stainless steel tank
  NEG strips in pump ports
  diagonal virtual area in one pump port (cross section between port and vessel) for determination of pumping speeds
three possible pump variations:
  NEG pumps hydrogen with aNEG between 0.5% and 3.5% (2.9% expected)
  TMPs for hydrogen or radon with their respective aTMP
  baffles with abaffle between 0% and 100% for radon
aims:
  find correlations between abaffle , aNEG and pressure ratios
  simulation of effective pumping speed of NEG, TMPs and baffles
  comparison with experimental ratios  effective pumping speed
  simulate radon suppression factor
Simulations of the Main Spectrometer

## Slide 67
Source: S. Görhardt: Background Reduction Methods and Vacuum Technology at the KATRIN spectrometers, PhD thesis, Karlsruhe 2014
Simulation of the Main Spectrometer
(MolFlow+)
main components:
  CF 200 ports on main vessel
  Baffles
  NEG strips
  Vacuum gauges
  TMPs

## Slide 68: Simulation of an effective pumping speed
Simulate pump as surface with an adsorption probability a
Determine pumping probability: w = Nads/Ndes
Calculate the effective pumping speed:
: mean molecular speed for mass M
Aport : desorption area (virtual area)
Nads : number of adsorptions in pump
Ndes : total desorption number
Aport
simplified model
pump port

## Slide 69
Determine pumping speed of TMP for mass of gas particle (Malyshev model)
Simulate pumping probability w = Nads/Ndes
Effective pumping speed:
TMP simulation
N2(28): 2650 ℓ/s
H2(2): 2060 ℓ/s
Rn(220): 1530 ℓ/s
Malyshev model:
O. Malyshev, Vacuum 81 (2007) 752–758
Pumping speed of TMP:

## Slide 70
NEG-pump simulation (without baffle)

## Slide 71
NEG-pump simulation (with baffle)

## Slide 72: Hydrogen outgassing and pressure at 20°C
NEG activation
desorption of H2
pressure at 20°C
F9 offset
Fit of pP3 versus pF9
  NEG pumping speed from pP3/pF9: 290 m3/s (a = 1.1%)
  Offset of Extractor gauge F9: 1.8·10-10 mbar
Fit of versus 1/T
  Desorption enthalpy of H2 on st. steel: 53 kJ/mol = 0.55 eV/H2
  Extrapolated pressure at 20°C: 2.6·10-11 mbar (gas corr. H2: 5.7·10-11 mbar)
Outgassing rate 1.4·10-12 mbar·ℓ/s·cm2

## Slide 73
Si-PIN diode
detection of transmitted β’s (mHz to kHz)
low background for T2 endpoint investigation
high energy resolution: ΔE = 1.48(1) keV (FWHM) at 18.6 keV
12 rings with 30° segmentation + 4-fold center = 148 pixels
  minimize bg, investigate systematic effects
  compensate field inhomogeneities of spectrometer’s analyzing plane.
KATRIN Main Detector