## Slide 1
Experience from 10 years of operating accelerators for ion-beam therapy at HIT: What are the key requirements for medical facilities?
12th December 2018
Workshop on Accelerator Performance and Concepts (APEC)
Christian Schömers
R. Cee, M. Ellerbrock, M. Galonska, A. Peters, C. Krantz

## Slide 2: What are medical facilities? In the accelerator context
Particle Therapy Centers doing patient treatment with Protons or heavier ions
66 Protons
11 Carbon (+ p)
ptcog.ch/ (Nov. 2018) Facilities in operation only.
28
20
17

## Slide 3: Rationales for ion beam therapy
Depth-dose profile: Bethe-Bloch
Range defined by starting energy (Bragg peak)
Less lateral scattering
Higher biological effectiveness
Possibility of lateral guidance (scanning)
Krämer & Durante, Eur. Phys. J. D 60 (2010)
Amaldi & Kraft, Rep. Prog. Phys. 68 (2005)

## Slide 4: How we do therapy State of the art: raster-scanning
Pencil beam is
guided by
  scanning magnets
monitored by
  multi-wire chambers
  ionization chambers
Spread out Bragg Peak
Accelerator delivers a spectrum of ion energies
courtesy of Siemens Health Care

## Slide 5: Example of a medical facility: Heidelberg Ion-Beam Therapy Center HIT
More than 5200 patients treated since Nov. 2009
Up to 71 patients treated per day
Treatment performed every day when scheduled
HIT
National Center for Tumor Diseases
Pediatric clinic
Gynecology
Surgery
Internal medicine
Radio- oncology

## Slide 6: Key requirements Reliability and Availability
Provide particle beams in therapy quality 335 days / year
Avoid any interruptions to guarantee therapy success
Beam precision and stability
Overview:
Facility design
Data supply
Operating concept
Maintenance concept

## Slide 7: Facility design Accelerator layout
Treatment room
Synchrotron,
(48 - 430 MeV/u)
Beam transport Beam shaping
Abort systems
Sources
Injector LINAC

## Slide 8: Facility design Basic strategy
The dose delivery must be:
Safe and precise
  Redundant beam abort systems
  Stable beam parameters (e.g. excellent power supplies)
Failsafe
  Redundancy of important components
    Sync-RF
    ECR Ion Sources
  Reliable components
    Warm iron magnets, not superconducting (decision of ~1998)
    Moderate field strengths, ~ 1,5 T
Fast: reduction of cycle time only without impact on 1.+2.
  Implementation of magnetic field control (2012)
  Implementation of dynamic intensity control (2014)

## Slide 9: Data supply Beam parameter
Precise dose delivery requires:
Large variety of beam parameters
  4 ion types (H, C, He, O)
  255 energies (48 – 430 MeV/u)
  4 spot sizes (4 – 13 mm)
  Intensity spectra (107 – 1010 part./ spill)
  36 gantry angles (0° – 360°)
A control system, that makes any combination available cycle-by-cycle
beam

## Slide 10: Data supply Data management
10 000 combinations per ion and target require intelligent...
Storage
  Stored in a database and front end controllers
  Therapy data stored in non-volatile flash memories
  Change of therapy data must be acknowledged (checksum!)
  Elsewise irradiation in treatment mode not executable
Check and adjustment procedures:
  Only ~2 % of these combinations, covering the whole range
  Daily accelerator performance check
  Quality assurance by medical physics team
Interpolation algorithm to fill entire parameter space

## Slide 11: Operating concept Typical therapy day
6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
Accelerator QA (1h):
  Measurements along the accelerator
  Beam size and position
  Particle rates and spill structure
  Beam currents in LEBT-LINAC-Section
  Output energy of LINAC
Patient treatment
Med. Phys.QA
QA,
Plan Verification
AccQA
R & D
or
Beam adjustment

## Slide 12: Operating concept Typical therapy day
6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
Medical Physics QA (2h):
  Room certification
  Position and width
  Beam abort systems
  Patient table movement
  Treatment plan verification
Patient treatment
Med. Phys.QA
QA,
Plan Verification
AccQA
R & D
or
Beam adjustment

## Slide 13: Operating concept Typical therapy day
6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
Treatment
  Up to 71 patients per day
  Coordinate 3 room operation
    Anesthesia
    Complex positioning
  19 min./day downtime in av.2018 (facility OR single room)
Patient treatment
Med. Phys.QA
QA,
Plan Verification
AccQA
R & D
or
Beam adjustment
H1
H2
Gantry
Room activity (15th Nov. 2018) 71 patients

## Slide 14: Operating concept Typical therapy day
6h 8h 10h 12h 14h 16h 18h 20h 22h 24h 2h 4h
R & D
  Bio-physics experiments
  Detector development (on-line beam monitoring)
  Accelerator R & D
Accelerator beam adjustment
Patient treatment
Med. Phys.QA
QA,
Plan Verification
AccQA
R & D
or
Beam adjustment

## Slide 15: Operating concept People
Accelerator team: ~20 people
  Physicists, engineers, technicians
  Maintenance/beam adjustment and operating is done by the same people
Availability of personnel:
  2 accelerator operators on-site, 24/7
  On-call duty for each of the 3 maintenance groups
  2nd level support (accelerator control system, Gantry)
In case of failure:
  Fast decision necessary, if therapy can restart within minutes
  Experienced personnel needed

## Slide 16: Maintenance concept Shutdowns
1 long shutdown
  Christmas – 1st week in January
  Time consuming issues
  Bad for patient management
5 maintenance blocks
  Maintenance Thursday + Friday
  Recommissioning and QA Saturday + Sunday
  Patients can continue treatment
12 single maintenance days
  Maintenance Monday morning
  QA in the afternoon

## Slide 17: Maintenance concept
Preventive maintenance
Parallel maintenance, tight schedule:
  HIT-internal groups (ion-sources, beam diagnostics, RF, ...)
  Siemens customer service for the beam application system
  External companies
    Infrastructure; cooling system
    Control system updates
    Gantry structure mechanical maintenance
   (Spatial) Coordination of different maintenance groups
No failure of a complete day since 2009
  Maximum breakdown: 5h

## Slide 18: Conclusion
Major criteria for a successful treatment
  Safety
  Accuracy
  Reliability
  Beam stability
Ways to get there:
  Use well known technology
  Redundancy
  Regular checks of beam parameters
  Experienced personnel
  Preventive maintenance

## Slide 19
Thank you!
And the entire HIT team, especially: R. Cee, M. Ellerbrock, M. Galonska, A. Peters, C. Krantz (MIT)