## Slide 1: SIS 18 VHF Cavity Status
JSt

## Slide 2: RF Tests have started
Expected:
  Resonanzfrequenz f0 82,4 MHz
  Shuntimpedanz RSK 4,6 MOhm
  Unbelastete Güte Q0 29.500
Measured before new contacting
  Resonanzfrequenz f0 82,4MHz
  Shuntimpedanz RSK 78kOhm
  Unbelastete Güte Q0 500
Measured with new contacting
  Resonanzfrequenz f0 82,4MHz
  Shuntimpedanz RSK 94kOhm
  Unbelastete Güte Q0 600

## Slide 3: Interpretaion, migration
Possible cause:
The dielectric loss faktor of the gap ceramics limits the shunt impedance to 735 kOhm.
Simulations show, that this cannot be the only effect
Conclusion: Expected metallic contermination in islands on the gap ceramics -> Cleaning!

## Slide 4: tentative ! measures
The RRF team is mostly on leavve or in Corvid quarantaine, so the expert meeting had to be rescheduled
With the present cleande gap 22-25 kV gap voltage should be possible. Enough for proof of principle (see: „Status of simulation studies on bunched beam slow extraction with very high frequency rf cavity by beginning of April 2020”, S. Sorge, P. Forck and R. Singh)
Later change gap material to lower dielectric loss factor (saphhire)

## Slide 5: Elektric field at 82,8 MHz