## Slide 1: BINP
6th BINP-FAIR Collaboration Coordination Workshop

## Slide 2: Branch Box (Conceptual design)

## Slide 3: Superconducting FRagment Separator (Super-FRS) Branch Box.
Boxing position: the responsibility of BINP is marked in yellow.

## Slide 4: Conceptual 3D model of the Branch Box and its cryogenic Transfer Lines.
BINP responsibility
The scope of cooperation of this work includes the technological design of the branch box (BB), which distribute helium supply to four SFRS experimental branches, Sub-cooler of the BB and four Transfer Lines for M, R, H, and L branches.

## Slide 5: Flowscheme of the Branch Box (January 2021).

## Slide 6: Interface of the Branch Box.
The interface was agreed with the GSI in September 2020.

## Slide 7: 3D model of the Branch Box.
3D model of the Branch Box with Warm piping and Pressure transmitters.
Vent Line
Pumping & Purging
High Pressure Supply
4.5K supply
5K return
50K supply
80K return
CWU supply
CWU return
MPL

## Slide 8: 3D model of the Branch Box.
3D model of the Branch Box with Warm piping and Pressure transmitters.
Pressure transmitters (34 pcs.)
Safety Tube

## Slide 9: Dimensions of the Branch Box.
Outer diameter is 2060 mm (2220 mm flanges), length is 5500 mm. The weight is ~14t
Branch L
Branch H
Branch R
Branch M
MPL
Multi-Purpose (cold-warm He) Line
CWU
Cooldown&Warmup Unit

## Slide 10: 3D model of the Branch Box.
Dimensions of the Branch Box wit Warm piping: 3 м width; 2,5 м height up to vacuum barriers branch M, R, H, L; 4 m height up to vacuum barrier CWU.
«Cold» Valves
«Warm» Valves
Safety valves
Safety Tube
Cryostat
Shield
Sub-cooler
Pressure transmitters

## Slide 11: Cryostat of the Branch Box
Outer diameter of the shell is 2060 mm, thickness is 12 mm, thickness of the cover is 45 mm, length of the Cryostat is 5500 mm. Material is St.St. The weight is ~7,3 t.
Cup
Flange
Leg
Shell
Cover
Pumping
Safety Tube
Vacuum barrier-CWU

## Slide 12: Shield of the Branch Box
Outer diameter is 1956 mm, length is 5340 mm, thickness of the shell is 4 mm. Material is Aluminum alloy. The weight is ~0,54 t. The shield is cooled using thermal bridges.
The shield is positioned into the cryostat using ball bearings with G11 supports.
Bottom Shield
Cover
Pipeline
Thermal bridges
Support
Flange
Top Shield

## Slide 13: Sub-cooler of the Branch Box
Pipes
Sub-cooler
Support
Outer diameter is 711 mm, length is 3550 mm (3150 mm cylindrical part), thickness of the shell is 8 mm, Material is St.St. The weight is ~0,56 t. Sub-coolers pipes: d=12x1 (DN10), l=9 m, 8 pieces.

## Slide 14
Sub-cooler of the Branch Box
Сharacteristic:
- The sub-cooler vessel volume – 1000 liters
The maximum design pressure is 20 bar against vacuum at all possible operating temperatures;
Temperature inlet - 6K;
Temperature outlet - 4,6K;
Pressure inlet - 4,0 bar;
Pressure outlet - 3,96 bar;
Flow rate - 174 g/s;
8 copper pipes OD12 x 1 mm (wall thickness) and length of the each pipe is 9 meters.

## Slide 15: Sub-cooler of the Branch Box
The maximum LHe level
The minimum LHe level
The LHe volume is 1000 liters in the sub-cooler vessel. 254 liters shall be guaranteed.
The maximum LHe level in vessel shall be less than 530 mm from the bottom, the minimum LHe level in vessel shall be more than 380 mm from the bottom.
The installation height of the Sub-cooler has to below 300 mm from the bottom (really - 243 mm).

## Slide 16: Cryostat, Shield and Sub-cooler of the Branch Box
Cryostat
Shield
Sub-cooler
Location of the Sub-cooler into the Cryostat

## Slide 17: Pipelines of the Branch Box.
4.5K supply
5K return
50K supply
80K return
Sub-cooler

## Slide 18: Pipelines of the Branch Box.
4.5K supply
5K return
50K supply
80K return
CWU supply
CWU return
MPL
Sub-cooler
Shield pipeline

## Slide 19: 3D model of the Branch Box.
3D model of the Branch Box with Working platform (6,4 m x 3,9 m; h=1,8 m).
Working Platform
Branch Box

## Slide 20: Conceptual design of the Transfer lines for M, R, H, and L branches.
TL03
TL04
TL05
Branch Box
TL06
Vacuum shell DN200
Al Thermal shield
Ball joint
G10 spacer
He pipe
Metal hoses

## Slide 21: Installation of the Branch Box.

## Slide 22: Plans:
After the final calculation of the flows, the diameters of the internal pipelines BB will be corrected.
The final calculation of the pipelines using the Caesar-II program
will be made.
The conceptual design of the Branch Box and four branch transfer lines will be prepared in September 2021.
The final design review will be prepared in January 2022.

## Slide 23: Thank you for your attention!