## Slide 1
MRPC design review meeting
Introduction to this meeting, to CBM TOF and observed aging effects
Ingo Deppner
Physikalisches Institut, Heidelberg University
23.11.2021

## Slide 2: Outline
A few introductory words
Introduction to CBM TOF
TOF FAIR Phase 0 program  eTOF at STAR/BNL  mTOF at mCBM at SIS18/GSI
- Beam test results
- Aging of MRPC gas and implications
STOP HERE -> See other presentations
Conclusions for the CBM TOF gas system
Summary and time line
2

## Slide 3: Introduction
External Reviewers:
- Marcello Abbrescia, CMS and EEE (Bari)
- Paulo Fonte, HADES (Coimbra)
What do we expect from the reviewers?
- Recommendations regarding R&D effort, new MRPC designs, project time line
3
Joint FAIR ECE 15 and ECSG 06 meeting (03.11.2021)
Joint FAIR ECE 14 and ECSG 05 meeting (31.05.2021)

## Slide 4: Introduction CBM
Compressed Baryonic Matter (CBM) Experiment
TOF
MVD: Micro Vertex Detector*
STS: Silicon Tracking System*
* inside magnetic field
MuCh or RICH
Muon Chamber System /
Ring Imaging Cherenkov Detector
TRD: Transition Radiation Detector
ToF: Time-of-Flight Detector
PSD: Projectile Spectator Detector
TRD
MuCH
MVD + STS
Dipole magnet
PSD
RICH
Tracking acceptance: 2.5o < θLab < 25o
Peak Rint is 10 MHz for Au+Au
Fast & radiation hard detectors
Free-streaming DAQ
4D tracking (space, time)
Online event selection and reconstruction
Data rate: 1 TB/sec
CBM Collaboration, EPJA 53 3 (2017) 60
T.Galatyuk, NPA982 (2019), update (2021)
4
Charged hadron identification is provided by Time-of-Flight (TOF) measurement

## Slide 5: Introduction CBM TOF
CBM-TOF Requirements
Full system time resolution sT ~ 80 ps
Efficiency > 95 %
Rate capability  150 kHz/cm2
Polar angular range 2.5° – 25°
Active area of 120 m2
Occupancy < 5 %
Low power electronics
(~100.000 channels)
Free streaming data acquisition
Multi-gap Resistive Plate Chambers (MRPC) are the most suitable TOF detectors fulfilling our requirements
5
FLUKA simulation: Au + Au collisions at Ekin = 11 AGeV, 107 interactions
X [cm]
Y [cm]
Anna Senger
Flux [Hz/cm2]

## Slide 6
Low resistivity glass
Introduction CBM TOF
6
thin float glass,   1012  cm
low resistivity glass,   1010  cm
ceramic   109  cm
13.5 m
8.75 m
Ceramics
Active area  118 m2

## Slide 7
Full size counter with close to final design for all regions build and tested
M4 and M6 full size modules constructed and installed at mCBM
 230 + 20* modules  1400 MRPCs  90000 channels
M6 Module (HD)
Low resistivity glass
MRPC2 (Tsinghua)
MRPC3/4 (USTC)
MRPC1a and 1b (NIPNE Bucharest)
M4 Module (HD)
BFTC* (ITEP Moscow)
BFTC* - Beam Fragmentation T0 Counter
Introduction CBM TOF
7
thin float glass,   1012  cm
low resistivity glass,   1010  cm
ceramic   109  cm

## Slide 8: CBM TOF MRPC key parameters
8
 | MRPC1a | MRPC1b | MRPC1c | MRPC2 | MRPC3 | MRPC4
# Gaps | 2 x 5 | 2 x 5 | 2 x 5 | 2 x 4 | 2 x 5 | 2 x 5
Gap size | 200 m | 200 m | 200 m | 250 m | 230 m | 230 m
Glass type | Low res. | Low res. | Low res. | Low res. | Float | Float
Glass thickness | 700 m | 700 m | 700 m | 700 m | 230 m | 230 m
MRPC size | 30 cm x 6 cm | 30 cm x 10 cm | 30 cm x 20 cm | 32 cm x 27 cm | 32 cm x 27 cm | 32 cm x 53 cm
# Strips | 32 | 32 | 32 | 32 | 32 | 32
Strip length | 6 cm | 10 cm | 20 cm | 27 cm | 27 cm | 53
Pitch | 0.90 cm | 0.90 cm | 0.90 cm | 1.0 cm | 1.0 cm | 1.0 cm
Impedance | 100  | 100  | 100  | 50  | 50  | 50 
Developed at | IFIN-HH (RO) | IFIN-HH (RO) | IFIN-HH (RO) | Tsinghua (CH) | USTC (CH) | USTC (CH)
Initially planed gas mixture: Tetrafluorethane/iso-Butan/SF6: 90%/5%/5%
Currently planed gas mixture: Tetrafluorethane/SF6: 97.5%/2.5%

## Slide 9: FAIR phase 0
FAIR Phase 0 programs of CBM-TOF
eTOF project at STAR@BNL (6912 channels) for long term stability test and physics results purpose
mTOF project at mCBM@SIS18 (1600 channels) for high rate and system integration test purpose
9

## Slide 10: eTOF status
eTOF successfully operated in RUN 19/20/21 (BES II)
arXiv: 1609.05102
10
Matching efficiency with TPC
eTOF PID
Matching efficiency with TPC dropped by 15%
Observation: FEE damages
TSU counter
USTC counter
Damaged channels
Damage on 30% of PADI channels
A change of protection diodes foreseen

## Slide 11: mCBM @ SIS18
11
mTOF
mTRD
mMuCH
mSTS
mPSD
mRICH
real Pb on Au collision
FAIR Phase 0: mCBM setup @ SIS18
mCBM is a full system test setup installed at SIS18/GSI dedicated for high rate detector and readout test including free streaming data acquisition and online event selection
Charged particle fluxes of up to 30 kHz/cm2
Target chamber
Simulation: Au+Au @ 1.24 GeV mbias
N. Herrmann

## Slide 12: Test setup in 2021
13
Beam
M4_5
911
910
(sealed) 901
900
M4_3
Target Au
Diamond T0
12
Triple stack
Double stack+ test counter
Gas guided MRPC3
25
Sealed MRPC2
Gas inlet of sealed MRPC
M4_5
MRPC3, 911 (USTC, float)
MRPC3, 910 (USTC, float)
MRPC2, 901 (sealed) (TSU)
MRPC2, 900 (TSU)
M4_3
MRPC1, 600 (Buc.)
MRPC1, 601 (Buc.)
12
M6_1 Module
55 cm
New inner wall FEE tested
MRPC4
Destructive test

## Slide 13: HV current example of 14 – 15 March
mTOF M4_4
mTOF M4_5
911
3(4) different intensities
2-3 times CBM conditions at highest IR
5 times CBM conditions for float glass counter
Test counter reached max current of the power supply
15 kHz/cm2
0.5 kHz/cm2
2.5 kHz/cm2
25 kHz/cm2
MVD+STS
100
13

## Slide 14: mCBM beam time results
Run 1459
Run 1459
MRPC3 (thin float glass counter)
14
Data loss issues (faced during beam time)
Bandwidth limitations
Get4 dropouts
Time slice losses
Next steps
Data analysis ongoing (much more to come)
Data losses have to be solved prior to 2022 beam time
More detector test at mCBM required
MRPC3 (thin float glass counter)

## Slide 15: Counter dark rates
Dark rates measured in 2020
About 1.5 h
Date - Time
Dark rate [Hz/cm2]
Gas pollution effect, observed in 2020, could be reproduced in 2021
Dark rate goes in saturation
Position (=> rate) ordering observed
15

## Slide 16: Gas aging & pollution
Aging test in Bucharest
Observations @ mCBM 2020
Gas pollution effect observed at mCBM at high rate (about 5 – 10 kHz/cm2)
Gas pollution effect was reproduced at IRSAM (Bucharest) with high gamma flux
X-Ray test at Beijing confirmed the gas pollution effect
The effect was minimized by sealing the MRPC
16
See CBM Progress Report 2020

## Slide 17: Gas aging & pollution
Observations: rapid increase of dark rate
Observed first @ mCBM 2020
Confirmed @ mCBM 2021
Mitigation strategy:
MRPC with clever gas management required (1 sealing and/or 2) forced gas flow through gaps)
Only parallel fishing lines proposed
=> new R&D round and beam tests required
New prototypes from the entire MRPC family will be tested at mCBM
First prototypes are being tested in cosmic test
17
Gas flow
Gas inlet
Gas outlet
Expected ‘ideal’
MRPC3/4
MRPC2
MRPC1

## Slide 18: Gas aging & pollution
Traces of NaF was found on the glass surface
Electrical field simulations performed
Mitigation strategy
Search for rectangular spacers
Increase gas flow
=> new R&D round and beam tests required
New prototypes will be tested at mCBM
Counter PRR shifted to 5/22 after mCBM test
18
Observations: continuous increase in dark rate
spacer
glass plate
glass plate
Electrical field simulations

## Slide 19
fresh counter gas
N2
circulation pump; ? l/min
filter
recycling station
gas flow in one modul purge mode 200 ml/min
cycling mode ? l/min
gas flow in all moduls purge mode 46 l/min cycling mode ? l/min
230 modules Total gas volume 25 m3
Roberto Guida (CERN)
Recuperation system
19
Stay with R134a (enhanced F-ion production for HFO in high rate environment)
Abandon iso-Butan (Aging)
Reduce fraction of SF6 to 2.5% (reduction of GWP, difficult to recycle)
Increase the flow rate (gas exchange in gaps within ? s)
Use clever gas management on MRPC (sealing, forced gas flow through the gaps)
Go for a recuperation system (reuse of gas, cost reduction, GWP reduction)
Recuperation system has to be developed
Conclusions for the TOF gas system

## Slide 20: Conclusion and Time line
High rate (mCBM) and long term stability (eTOF) test performed - float glass counter performed beyond specs
Damage on preampl. observed at STAR, could not be reproduced with highest intensity beam and distructive test at mCBM
MRPC aging and gas pollution effects at high rates observed - mitigation strategies developed => more R&D required => one more beam test needed
Sealed and gas guided MRPC designs presented – Recommendations highly welcome
Conclusions for the TOF gas system discussed
Preparation for the upcoming beam time ongoing (counter development)
Layout of gas system with recuperation drafted (First discussions with ICSI/RO)
20

## Slide 21: Conclusion and Time line
Time line and major milestones
MRPC design review meeting: 23.11.2021 (today)
ASIC and Electronics PRRs: Q4/2021
mCBM beam time: 02/22 – 04/22
Counter PRR Q2/2022
Counter production start: Q3/2022
Module production start: Q4/2022
TOF ready for beam: end of 2024
21

## Slide 22: Thank you for your attention
Contributing institutions:
Tsinghua Beijing,
NIPNE Bucharest,
GSI Darmstadt,
TU Darmstadt,
USTC Hefei,
PI Heidelberg,
ITEP Moscow,
HZDR Rossendorf,
CCNU Wuhan,

## Slide 23: Alternative gas search in CBM TOF
(a) Pure HFO (b) HFO/i-C4H10/SF6 90/5/5 (c) HFO/SF6 95/5 (d) HFO/i-C4H10 97/3 (e) HFO/i-C4H10 95/5
(RPC2020)
Alternative gas mixtures were investigated
Time resolutions in the order of 80 ps to 100 ps were obtained
Gas mixtures with HFO fulfil our TOF requirements
Working point is shifted by about 2000 V
Measurements of the eff. Townsend coefficient show no major difference for gas mixtures with more than 2 % SF6
The instability of HFO in comparison to R134a is counterproductive in a high rate environment
Xingming at 16:35
Wang Botan, Wang Yi
18

## Slide 24: Gas aging & pollution
Aging test in Bucharest
Observations: rapid increase of dark rate
Observed first @ mCBM 2020
Confirmed @ mCBM 2021
Gas pollution effect observed at mCBM at high rate (about 5 – 10 kHz/cm2)
Gas pollution effect was reproduced at IRSAM (Bucharest) with high gamma flux
X-Ray test at Beijing confirmed the gas pollution effect
The effect was minimized by sealing the MRPC => See Botans talk
MRPC with clever gas management required (forced gas flow through gaps)
=> new R&D round and beam tests required
See CBM Progress Report 2020
13
Observations: continuous increase in dark rate
See Marianas talk
spacer
glass plate
glass plate

## Slide 25: Gas aging & pollution
Mihai Petrovici
Mihai Petrovici
Aging test in Bucharest
Aging test in Bucharest
22
Aging test performed at IRASM (360 kCi)
Glass inspections performed with various methods (SEM, XPS, AFM RBS, non-RBS, THz-TDS)
Glass resistivity remains unchanged

## Slide 26: Inspection of low resistivity glass
23
F. Carnesecchi et al. NIMA A 908 (2018) 285-290

## Slide 27: Inspection of float glass
10 m
28 m
Unused float glass
24
Yongjie Sun, Kaiyang Wang, Xinjian Wang

## Slide 28: eTOF introduction
eTOF successfully operated in RUN 19/20/21 (BES II)
Fixed target mode
arXiv: 1609.05102
Gold target
25
sNN /GeV | # coll. Events | Year
7.7 | 100 M | 2021
9.1 | 150 M | 2020
11.5 | 230 M | 2020
14.6 | 320 M | 2019
19.6 | 580 M | 2019
Collider mode
sNN /GeV | # coll. Events | Year
3.0 | 2 B | 2021
3.5 | 100 M | 2020
3.9 | 50 M | 2020
4.5 | 100 M | 2020
5.2 | 100 M | 2020
6.2 | 100 M | 2020
7.7 | 50 M | 2020
9.2 | 50 M | 2021
11.5 | 50 M | 2021
13.7 | 50 M | 2021
Fixed target mode

## Slide 29: eTOF status
Preliminary
System synchronization on the level of 35 ps
System time resolution better than 85 ps
PID capability demonstrated
Physics analysis started
Fixed Target s = 7.7 GeV
26
Matching efficiency with TPC
Timing resolution
eTOF PID

## Slide 30: FEE: PADI damages at BNL
-615mV
Chip 1
Chip 2
Chip 3
Chip 4
Ch1
Ch1
Ch1
Ch1
Chart: PADI board #2R of module 23-3
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32
positive input to GRD | 325 | 325 | 324 | 324 | 323 | 324 | 324 | 324 | 324 | 324 | 323 | 324 | 324 | 324 | 324 | 323 | 323 | 323 | 323 | 323 | 323 | 323 | 324 | 322 | 322 | 322 | 322 | 323 | 323 | 322 | 322 | 322
negative input to GRD | 0.163 | 0.152 | 0.18 | 0.143 | 0.171 | 0.506 | 0.175 | 0.14 | 324 | 0.187 | 0.167 | 0.112 | 0.204 | 0.35 | 0.201 | 0.5 | 323 | 0.162 | 0.191 | 0.216 | 0.206 | 0.98 | 0.227 | 0.42 | 322 | 0.217 | 0.216 | 0.164 | 0.237 | 0.248 | 0.4 | 322
output to GND | 1.34 | 1.36 | 1.388 | 1.35 | 1.38 | 1.821 | 1.37 | 1.34 | 1.4 | 1.38 | 1.32 | 1.42 | 1.57 | 1.4 | 1.79 | 1.39 | 1.42 | 1.45 | 1.44 | 2.21 | 1.46 | 1.67 | 1.44 | 1.45 | 1.39 | 1.47 | 1.48 | 1.64
50 
50 
2.2 nF
2.2 nF



Matching efficiency with TPC dropped by 15%
Damaged channels
27
Observation
Threshold setting for individual channels not possible after damage
-200mV
threshold
Negative input resistors lose nom. value
Damage on 30% of PADI channels
Reason for damage not fully clear
Characteristic of damage understood
Damage could not be provoked in mCBM
A change of protection diodes foreseen
Not a likely risk for CBM
Protection diodes
TSU counter
USTC counter
Observation