## Slide 1
Alberto Pullia
INFN - Milano University of Milano
Department of Physics
12th AGATA Week June 11-13, 2012 GSI Germany
June 13, 2012
“DIGI-OPT12” digitizer for AGATA and GALILEO*
*Technical details in white paper: “DIGI-OPT12: 12-channel 14/16-bit 100/125-MS/s Digitizer with Optical Output for AGATA/GALILEO” version 1.5

## Slide 2
Mi-Pd group technical meetings from late 2010 to now
Oct. 6, 2010 - Legnaro
Dec. 15, 2010 - Padova
Febr. 1, 2011 - Legnaro
March 1 2011 - Padova
Apr. 21, 2011 - Padova
May 5, 2011 - Legnaro
June 21 2011 - Padova
July 19, 2011 - Legnaro
Oct. 13, 2011 - Legnaro
Dec. 21, 2011 - Legnaro
Jan. 17, 2012 - Legnaro
Apr. 27, 2012 - Legnaro
May 16, 2012 - Padova

## Slide 3
System parts and connections (1 AGATA crystal)
Spare
MDR cables
Pre-processing Card
Padova
GTS
link
PC + PCIe expansion box
fibers
Clock-Distribution
and Control Unit
Padova
3 Segment ADC Modules
(12 channels each)
Milano
Core ADC
Module
(same HW)
Milano

## Slide 4
Backplane-ready design
Backplane for clock, control signals and PS
Backplane
connectors
Conceptual design
Use of a backplane is foreseen in AGATA, which greatly reduces cable burden
MDR connectors
Real thing
MDR connectors

## Slide 5
Optical
transmitter
ADC card
(this presentation)
MDR Connector #1
6 segments or 1 core
MDR Connector #2
6 segments or 1 core
TE 6469028-1
and/or
mini-HDMI
clock
Slow Control
Analog
Signal
Conditioner
ADCs
I2C devices
Slow control unit
Power Supply
conditioner
SPI devices
Snap12
clock, sync, test
i2c, spi
+2.0V
+3.3V
TE 120955-1
and/or
screw connector
Low noise
Minimum power consumption
High integration
High flexibility
ADC with integrated JESD204A encoder/serializer
Key words:
FPGA-less design!

## Slide 6
Detailed schematic diagram of ADC card (ver 3.6)
Design completed in July 2011
Layout completed in January 2012

## Slide 7
DIGI-OPT12 card - ver 3.6
It’s just a little bit larger than a CD-ROM
The first card has been delivered on April 2012

## Slide 8
Size of a single ADC card
DIGI-OPT12

## Slide 9
Set of thre piled up cards
Backplane
connectors
MDR connectors
Backplane
connectors
MDR connectors
Front view
Rear view

## Slide 10
The card – top view
EMI filters
clock distr
analog front end
ADC, CK, TX manual reset
160 mm
120 mm
Power (3.3V, 2V) via backplane
ck, sync, test i2c, spi via backplane
All signals via mini-HDMI
MDR in
Power (3.3V, 2V) via screw header
SNAP12 TX (not connected)
conn’s for core/segm mezzanine
J5
J6
J7
ADC dual
quad digi-pot
Optional fast out with LLD
LEDs

## Slide 11
Principal properties of DIGI-OPT12 card
14 (or 16) bits, 100MS/s (up to 125) - coding (8b/10b) and SER integrated in ADC chip
Optical output signals (SNAP12, 1 fiber per analog channel)
Compact (120mm x 160mm) and low power (~10W including laser)
Wide choice of factory-set configurations including:
 differential (AGATA) or single-ended input
 segment or core configuration
 AC or DC coupling with or without active input stage (4 setups)
 preset value of common-mode DC voltage
 preset parameters of antialiasing filters
 on-board clock distribution with or without PLL conditioning
End-termination for both differential and common-mode signal components
Remotely-controlled adjustment of DC offset
Remotely-controlled selection of energy range (5 or 20 MeV)
Remotely-controlled setting of ADC and clock-distribution chip parameters
Multiplexed test-pattern input for time synchronization of ADCs
Optional interleaved mode (e.g. 6chs@200MS/s, 4chs@300MS/s, etc)

## Slide 12
Clocking the ADCs
Clock jitter is the key parameter to achieving a good SNR/ENOB in the digitized signal
tJ = rms jitter on sampling clock fA = highest analog frequency being digitized
Noise of on-board clock-distribution chip (AD9522-3 or pin-compatible AD9520-3):
Additive clock jitter: < 260fs
clock distribution only, PLL off
Absolute jitter 1: ~ 300fs
in 12kHz-20MHz BW, with clean clock input. PLL on, with loop BW of 55kHz.
Absolute jitter 2: ~ 400fs
in 12kHz-20MHz BW, with jittery clock input. PLL on, with loop BW of 2kHz.
Required clock jitter as a function of fA and ENOB
Use of a high-quality clock distribution chip is the key to get the lowest jitter. Our choice is Analog Devices AD9522-3, with an embedded PLL (on/off options).
fA in AGATA

## Slide 13
System architecture: input stage and offset circuit
Offset provided through an i2c digipot and opamps
1.5 V
3.0 V
Digipot 256 positions
i2c
DC - os
DC + os
(DC – os) / 2
(DC + os) / 2
DC + os
DC - os
from preamp
“20 Mev” range
“5 Mev” range
DC coupling, active input stage - default setup
to ADC
In –
In +

## Slide 14
System architecture: signal path
+
Offset
Regulation
Fast
Amplifier
ADC
Antialias
Laser
Divider
High
High
i2c
spi
Detector signal
Signal path for “20 MeV” range measurements (-25% DC offset required to get the full range)
Test Pattern
i2c I/O

## Slide 15
System architecture: signal path
+
Offset
Regulation
Fast
Amplifier
ADC
Antialias
Laser
Divider
High
Low
i2c
spi
Detector signal
Signal path for “5 MeV” range measurements
Test Pattern
i2c I/O

## Slide 16
System architecture: signal path
+
Offset
Regulation
Fast
Amplifier
ADC
Antialias
Test Pattern
Laser
Divider
Low
Low
i2c
spi
Detector signal
Signal path for time calibration/synchronization
Damiano alg.
i2c I/O

## Slide 17
Setup for testing the DIGI-OPT12 card
i2c from PC
power
DIGI-OPT12
clock
Notes: 17

## Slide 18
Heat dissipation
The card on the test bench with tiny heat sinks on the ADCs. Thermographic picture below
Fan-less
With USB powered fan for notebooks
Fischer tiny heat sink
CK
IN
i2c
2.0V, 3.3V
PLL lock LEDs
Notes: 18

## Slide 19
Clock signals captured at the ADC’s input pins
Clock signal @ 100 MHz is clean and symmetric !!
The clock signals are provided to the ADCs through the on board clock-distribution chip
Notes: 19

## Slide 20
Eye diagrams of the 12 high-frequency digital signals
Data stream = 2 Gbps per lane. Eye diagrams are nicely open !!
500 ps
The 12 encoded/serialized datastreams are provided to the optical TX at high frequency
Notes: 20

## Slide 21
Specifications
ADC | 14 or 16* bit, 100MS/s with JESD204A (8b/10b) and SER interface
Channels | 12 per card
Analog input | Differential, MDR connectors as specified in AGATA preamp white paper v. 3.2
Configurations | “Segment mode” (default, 12 chs) or “Core mode” by insertion of passive piggyback PCB (3 chs + spares)
Clock input | Differential LVPECL through mini-HDMI or backplane
Sync and test pattern input | Single-ended LVCMOS with static toggle through e-SATA connector
Control input | I2C and 3-wire SPI through mini-HDMI (HDMI type C) connector
Output | Optical, 1 fiber each digitized channel (see datasheet of ReflexPhotonics SN-T12-C00601 SNAP12)
Power Supply | 3.3V @ 2.5A and 2.0V @ 0.6A
Power cons | ~ 10W per card
Size of card | 120mm x 160mm
Range control | Remotely controlled range selection: (a) “20 MeV range” (with 25% offset displacement) and (b) “5 MeV range”
Offset control | Remotely controlled within +-30% of full swing
ADC param control | Remotely controlled full set of ADC and JESD204A parameters (see datasheet of NXP ADC1413D)
Clock param control | Remotely controlled full set of clock-distribution parameters, includind switching on the embedded PLL for zero-delay option (see datasheet of AD9522-3)
Pulser param control | Remotely controlled full set of pulser parameters in “Core mode” (see AGATA preamp white paper v. 3.2)
Options | Interleaved mode (e.g. 6 equivalent channels @ 200MS/s), single-ended analog inputs, built-in clock generation
See also white paper: “DIGI-OPT12: 12-channel 14/16-bit 100/125-MS/s Digitizer with Optical Output for AGATA/GALILEO” version 1.5 (or later)
*Pin-to-pin compatible ADC, mod. ADC1613D, is available with a maximum sampling frequency of 125 MHz

## Slide 22
Road map and conclusion
April - June 2011
ADC1413D semi-qualified @ Padova
Damiano algorithm for cancellation of random latencies of ADC/state machine
June 2011
Schematic diagram completed (in Orcad Capture)
July - September 2011
Translation of schematics in other CADs (Orcad ConceptHDL, Zuken Cadstar)
September - December 2011
Layout synthesys
Spring 2012
First prototypes ready for testing
Spring-summer-autumn 2012
Tests, qualification, preproduction for GALILEO

## Slide 23
ADC card mounting: fully stacked or paired-and-stacked
Paired and stacked
Fully stacked

## Slide 24
System architecture: the board
Analog in
Analog in
The path from the MDRs to the ADCs goes through a mezzanine card  easy segment or core configuration

## Slide 25
System architecture: the board
Analog in
Analog in
The path from the MDRs to the ADCs goes through a mezzanine card  easy segment or core configuration

## Slide 26
Clock-Distribution and Control Unit
ADC Units Clock & Slow Control
clock,
sync/test,
slow control
Notes: 26