## Slide 1: Position and beam size measurements for unbunched beams in transfer lines
Akihisa Toyoda (KEK/IPNS)
on behalf of the J-PARC hadron beamline group

## Slide 2: Outline
Overview of the J-PARC Hadron beamline and its monitors
Features of the Hadron beamline profile monitors
Performance of the Hadron beamline profile monitors
Summary

## Slide 3: J-PARC Hadron Experimental Facility
J-PARC Web page
MR 30 GeV
Hadron
Experimental
Facility
30 GeV proton beam from the Main Ring (MR) is slowly extracted to the Hadron (HD) Experimental Facility.
  Current operating cycle: 5.2 s
  Current spill length: 2 s
time
MR PS pattern
Inj.
Acc.
Ext.
RCS 3 GeV
Linac 0.4 GeV
2 s
5.2 s

## Slide 4: J-PARC Hadron Primary Beamline
A line
B line
C line
A line 30 GeV 65 kW (7.0 x 1013 protons/spill)
  3 experimental areas using secondary beam from the T1 target
B line 30 GeV 24 W (2.6 x 1010 protons/spill)
  Branching out part of A line beam.
  1 experimental area using primary proton beam directly
C line 8 GeV 0.33 kW (2.5 x 1012 protons/spill, 9.6 s cycle, 0.5 s spill length)
  Bunched slow extraction beam
  1 experimental area using secondary muon beam from the target
T1 target
Ext. point
A-B branch point
Dump
Vacuum foil

## Slide 5: Beam Diagnostic Devices in Primary Beamlines
Beam Loss Monitors
Profile Monitors (OTR: Optical Transition Radiation)
Profile Monitors (RGIPM: Residual Gas Ionization Profile Monitors)
Target Monitor
Spill Monitor
Beam Intensity Monitor (Scintillation)
Beam Intensity and Short Pulse Monitor (RGICM)
Beam Intensity Monitor (Ion Chamber)
Profile Monitor (Ion Chamber)
Low vacuum
(about 1 Pa)
High vacuum
(about 10-7 Pa)

## Slide 6: RGIPM: Residual Gas Ionization Profile Monitor
100
280
280
Type-100
Beam
Permanent Magnet
(Sr Ferrite)
0.3 Pa vacuum
Negative electrode
(-100 V)
Electric field shaping electrodes
32 channel readout electrodes
(1 mm pitch)
uniform magnetic field
(400 gauss)
Y. Sato
3rd IPM WS

## Slide 7: Features of Hadron Beam Profile Monitors(RGIPM)
Non-destructive
Maintenance free
  Use the Sr ferrite permanent magnet
  No amplification devices
  High radiation resistant materials: inorganic resistances, ceramic or glass epoxy substrate, and so on.
  Operating in high radiation environments such as 1 m upstream of A line target called T1(50 % loss)
    about 2 MGy/year (1014 protons/s, 2500 h/year operation)
  No maintenance and replacement since the first beam in 2009.
Shot by shot or time division (about 8 division) measurement
  Monitor beam width and beam position and stop abnormal beams
Different types of RGIPM depending on beam size
  Beam detection region: 100 mmf, 300 mmf, and 400 mmf
  Readout electrode pad pitch: 1 mm, 3 mm, 6 mm, 10 mm

## Slide 8: Using the RGIPM as the HD Profile Monitors
Solve small S/N ratio problem
  Slow extraction and single pass beamline: Signal per unit time is much smaller than that of pulsed and circulating beam Small S
  Large background: foil, target, dump, SX devices, and so on Large N
  MCP to enhance signal amount cannot be used: Large background causes large noise, and saturate MCP signal.
  Successfully enhance signal by controlling vacuum to about 1 Pa level
  Successfully improve S/N ratio by integrating the signal.
Suppress beam width broadening effect
  Relatively low vacuum causes the beam width broadening due to the diffusion of drifting electron.
  Successfully suppress this effect by applying a magnetic field about 200 to 400 Gauss depending on the beam size.

## Slide 9: Photos of the Installed RGIPM
Type-300
1140
Permanent Magnet
Horizontal
Vertical
Horizontal
Type-300 rear view of internal electrode
Vertical
readout
electrodes
No destructive material on beam

## Slide 10: Readout Electronics and DAQ
Beam extraction 2 s
0 s
5.2 s
time
charge
integration
A/D
conversion
off-beam
measurement
on-beam
measurement
0 s
5.2 s
time
off-beam
measurement
on-beam
measurement
background subtraction to enhance S/N
Time division profile is also in use by using multi-scanning
32 ch charge integrator board on VME
  Input: 1, 4, 40, 400 nC / 5 V
64 ch scanning A/D board on VME
EPICS IOC is running on the VME CPU board
  EPICS records are provided.
  EPICS channels are recorded by the EPICS Archiver Appliance

## Slide 11: Display Screen of the RGIPM
GUI using wxPython + pyepics
On-line gaussian fitting provides width, position, and max values.
Vertical
Horizontal
T1
Sigma
3.6 mm
Sigma
2.4 mm

## Slide 12: Emittance Measurement @A line 30 GeV beam
[mm]
[mm]
Beam optics calculated by the TRANSPORT
Beam emittance is evaluated by fitting the beam envelope.
[kW] | [yymmdd] | εH(2σ) | εV(2σ) | T1 σ (H [mm],V [mm])
64.5 | 210602 | 1.25 | 1.11 | 2.56 x 0.89
RGIPM position
RGIPM position

## Slide 13: Profile Drift in Typical One Spill@T1
X
Y
SX
SY
[mm]
[mm]
[mm]
[mm]
Time Division #
Time Division #
Profile drift can be observed on the RGIPM monitor.
X drifts about 2 mm, Y drifts about 0.3 mm.
SX is stable and SY is gradually narrowing.
Profile stability in one spill is sufficient to maintain target health.
Secondary beam quality (K/p ratio) is also stable enough for the experiments to use.
Mean X
Mean Y
Sigma X
Sigma Y
Spill Start
Spill End
Spill Start
Spill End

## Slide 14: Long-Term Profile Stability@T1
Mean X
Mean Y
Sigma X
Sigma Y
[mm]
[mm]
[mm]
[mm]
5 sigma level
Jun. 11, 2021
Jun. 30, 2021
Jun. 11, 2021
Jun. 30, 2021
5 sigma level
5 sigma level
5 sigma level
±0.63 mm
±0.11 mm
±0.19 mm
＋0.33 mm
The Y position which is the most important parameter for the target health and secondary beam quality has enough good long-term stability.
The other three parameters are also stable enough for the experiments.

## Slide 15: Long-Term Profile Yield Stability@T1
Normalized Profile Sum divided by Beam Intensity
+4% level
-4% level
Jun. 11, 2021
Jun. 30, 2021
Sum divided by beam intensity has the stability of about ±4% level.

## Slide 16: Long-Term Profile Stability for 8 GeV Bunched Beam @B line Collimator
Mean X
Mean Y
Sigma X
Sigma Y
[mm]
[mm]
[mm]
[mm]
Mar. 9, 2023
Mar. 9, 2023
Mar. 16, 2023
Mar. 16, 2023
5 sigma level
5 sigma level
5 sigma level
5 sigma level
±0.19 mm
±0.05 mm
±0.055 mm
±0.040 mm
Beam profile stability is better than that of the 30 GeV A line beam.
Stability of the beam profile is good enough for the experiments.

## Slide 17: Summary
Hadron RGIPMs are stably running for the 30 GeV 65 kW slow extraction beam and the 8 GeV bunched slow extraction beam.
  Short and long-term profile stability is good enough for the target health and secondary beam quality.
Maintenance free and no replacement since the first beam in 2009.
Operating even in high radiation environments such as 1 m upstream of the T1 target.
  No performance degradation is observed yet.

## Slide 18: Vacuum dependence
0.3
0.4
Normalized Profile Sum divided by Beam Intensity
Vacuum [Pa]
Clear dependence of the profile sum on the vacuum.
One of the reason why the profile sum trend fluctuates is due to the vacuum change.

## Slide 19: Vacuum dependence of profile width
Sigma X
Sigma Y
[mm]
[mm]
0.3
0.4
Vacuum [Pa]
0.3
0.4
Vacuum [Pa]
Profile width is less dependent on the vacuum compared to that of profile sum.
This indicates that the effect of profile broadening due to vacuum is sufficiently small compared to the profile width.

## Slide 20: Correlation between MR MWPM and T1 RGIPM
Sigma X
Sigma Y
Mean X
Mean Y
MR MWPM
MR MWPM
T1 RGIPM
T1 RGIPM
Small dependence of the profile position of the T1 RGIPM on that of the MR MWPM.
The dependence for the Sigma X is small, and the dependence for the Sigma Y is not clear.

## Slide 21: OTR(Optical Transition Radiation)
rotary drive
mechanism
OTR monitor is a profile monitor utilizing the transition radiation generated at the thin foil.
The foil is made of Ti (thickness 5 mm) and can be inserted or removed as needed by the drive mechanism.
Transition radiation light is focused by the optical system, amplified by the Image Intensifier (IIT), and measured by the radiation-tolerant camera.
The near maintenance-free operation is achieved by relocating equipment relatively sensitive to radiation approximately 5 m away from the beamline.

## Slide 22: HGRGPM (High Gain Residual Gas Profile Monitor)
S
N
Negative
electrons
Duct
Phosphor
Screen
Ground
Optics+IIT+camera
Beam
Observe the light from the phosphor using an optical system, IIT, and camera
Achieves signal amplification by converting the signal into light with less background and transmitting the light away from the beamline.
Improve S/N by preparing a thin phosphor(Gd2O2S(Tb), 20 mmt, no coating)
R&D work in progress
magnet
magnet
Optics

## Slide 23: Photos of the Installed RGIPM
Type-100
Type-300
900
1140
Vertical
Horizontal
Vertical
Horizontal

## Slide 24: T1 Target
Phys. Rev. Accel. Beams 25, 063001
beam
H. Watanabe, HPTW2023
Indirectly water-cooled fixed type target divided into upper and lower parts
Vertical beam position stability is important to suppress the target temperature rise

## Slide 25: Magnetic Pole Cross Section of the Lambertson Magnet
Magnetic field
To B line
To A line
Field free hole
Y. Komatsu et al.
Proceedings of the 17th Annual Meeting of Particle Accelerator Society in Japan
Yoke
The beam that passes through the field free hole goes straight to the A line, and the beam that passes through the magnetic field region is bended and heads towards the B line.
Vertical beam position stability is important for B line intensity stability.