## Slide 1: HITRAP Lessons from the beam time 2024 and outlook to 2025
Zoran Andelkovic
GSI / Accelerator Operations / Decelerator

## Slide 2: HITRAP facility overview
IH Structure
RF Quadrupole
4.03 MeV/u
4.03 + 0.495 MeV/u
4.03 + 0.495 + 0.006 MeV/u
Magnet +
0.1 mm
Slit
Detector
Detector
Detector
Magnet +
0.1 mm
Slit
BB
BB
Rep. Rate
45-60 sec.
Trap
Goal 105 slow HCI  trap
transport & deceleration
energy separation
trapping, cooling, extraction
Reached so far
in units of beamtime:
  beam transport to the trap
  trap operation test run E057
  extraction and mass separation
  E057 with U91+
  …
2024
2025
✔
✔
?
✔

## Slide 3: Shutdown work 2023-24
✔
Preparation for beam time ‘24
detector development
e-cooling development
CS expansion to GTR5
stabilization IH amplitude
EBIT
+ GTR7
GTR1-4
GTR5
GTR6
✔
✔
42mm

## Slide 4: Electron cooling development & Experiments
» like e- cooling in a storage ring at extremely low-E
» an experiment with constant research demand
(now with a generation change of PhD students, TU-DA)
» major progress, first worldwide:
ions electrons cooling online operation
20 kV
6 T
4 K
40 cm
e-
e-
HCI
✔
✔
✔
NE14 (ex-reinjection)
new passage through shielding
two new walls
New HYIG with P. Micke @ HITRAP
talk by T. Stöhlker on Thursday
laser lab construction between beam times
preparation for ion delivery from the trap

## Slide 5: Beamtime overview 2024
15.04 | 16.04 | 17.04 | 18.04 | 19.04 | 20.04 | 21.04 | 22.04 | 23.04 | 24.04 | 25.04 | 26.04 | 27.04 | 28.04
36-Ar 18+ ( m/q = 2.0 ) |  |  |  |  |  |  |  |  |  |  |  |  | 
Setup |  |  | parallel oper. Er |  | unstable |  | setup for parallel oper. |  |  |  |  |  | 
16.02 | 17.02 | 18.02 | 19.02 | 20.02 | 21.02 | 22.02 | 23.02 | 23.02 | 24.02 | 25.02 | 26.02 | 27.02
12-C 6+ ( m/q = 2.0 ) |  |  |  |  |  |  | 18-O 8+ ( m/q = 2.25 ) |  |  |  |  | 
Beam setup (UNILAC + SIS + ESR) |  |  |  |  |  |  | Setup |  |  |  |  | 
A total of ca. 4 days with a stable beam
A total of ca. 6 days with a stable beam
up to 2022: 16 tests (3-7 days, 45 sec / shot)
= ca. 1 (one) hour of operation @ 50 Hz
3 on-line tests since recommissioning
…
2022
2023
2024
2025
some problems due to setup & parallel operation
shift to 36Ar helped a lot
overall quote OK

## Slide 6: Excellent results 2024
RFQ + beam transport ion trap / B = 6 T
slow ions in front of the trap
a mix of fast components
6 - 0.5 MeV/u
SUCCESS!
worldwide first time HCI were stopped after production, estimate 104-105 particles
slow ions behind the trap
105 ions 36Ar18+
stopped, stored & extracted

## Slide 7: Observed problems
Chart: IH power curve
| 1 | 2 | 3 | 4 | 4.5 | 5 | 5.25 | 5.5 | 5.75 | 6 | 6.1 | 6.2 | 6.25 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.75 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.25 | 7.5 | 7.75
Year 2023 | 3.3 | 11.5 | 25 | 43.75 | 54.75 | 67.25 | 81.25 | 96.75 | 100.25 | 103 | 107 | 110.5 | 114 | 117.5 | 121.5 | 126 | 143 | 150 | 155.75 | 160.75
Year 2011 | 53 | 60 | 67 | 75 | 83 | 91 | 100 | 109 | 118 | 129 | 140 | 152
Feb. 2024
18O8+ 7.1 V
(m/q = 2.25)
Mar. 2010
86Kr33+; 7.85 V
(m/q = 2.61 = 238U91+)
36Ar
18O
238U
1
2
GTR3QT3 (internal IH triplet & steerer)
last shift ca. 200 liters/day
leak from the steerer subunit
intensity loss ca. 85% if switched off

## Slide 8: Consequences for the experiments
Achieved in this beam time:
beam transported to the trap
  Prerequisite for E057
ions stored in the trap
  Prerequisite for e- cooling
ions decelerated in the trap
  new HYIG & ERC Advanced Grant
Open issues:
expansion of the control system
  includes mostly FESA + independent operation
reestablish IH Amplitude 8 V
  several action items in the Shutdown Plan 2024
  invest in Solid State amplifiers (ca. 750 k€, 2 years)
  restriction to m/q < 2.20
water leak TR3KX4 (IH triplet & steerer)
  new unit (1 year & ca. 150k€) + mounting time
  temporarily decommission (ca. 85% intensity loss)
setup time and repetition rate
  several beam times per year are essential
  revisit the idea of a local source of 4 MeV/u
with the right conditions, some beams possible in 2025
78Kr
(m/q = 2.16)
238U
(m/q = 2.61)
2025

## Slide 9: Thank you for listening!
Also thanks to the DEC team and collaborators:
S. Fedotova, W. Geithner, F. Herfurth, J. Ködel, N. Kotovskiy, I. Kraus, K. Mohr, D. Neidherr, D. Racano, S. Rausch, N. Stallkamp, S. Trotsenko, G. Vorobyev, W. Nörtershäuser, D. Zisis

## Slide 10: End of beam time fun
so despite all the stress…
… always look at the bright side of life

## Slide 11: Perspectives
Engineering run 2023
  expand FAIR CS to GTR5
  reach RF amplitudes > 7V
shutdown work 2024
PhD students on cooling trap
further beamtime depends on the schedule and ESR availability
alternative scenarios investigated
  a study report available
a no-fail plan / units of beamtime:
  transport through GTR5
  online cooling; test run E134 with Ar18+
  extraction and mass separation
  E134 with U91+
  multiplexed operation
  single ions for HYIG (optical clocks)
  transport to the 2nd floor
  E130 with Bi82+
  …
current status: ca. 106 ions @ 6 keV/u
2024
2025
1, 4
2, 5
3, 6
7, 8

## Slide 12: Commissioning experience
reconstruction
(2013-14)
Why is a linear decelerator NOT simply a reversed accelerator?
the emittance grows
narrow acceptance
fast beams stay in
low repetition rate
2007
…
2014
…
2022
16 On-Line Tests (3-7 days, 45 sec / shot)
= 1 (one) hour of operation @ 50 Hz
45 sec delivery rate:
set one dipole = 20 min
set two dipoles = 20*20 min = 6h
360° RFQ rough scan = 3h
360° RFQ full scan = one shift
if everything OK:
» ca. 5 days for beam setup
» ca. 5 days for experiment
=> minimum two online beams per year for progress
One-shot energy analyzer
0.1 mm slit
Beam
0.5 T perm. Magnet
MCP
Mirror & CCD

## Slide 13
Ion | A/q | IH power / kW | IH Tank / V (extrap.) | RFQ power / kW | RFQ Tank / V (extrap.)
14N7+ | 2.00 | 100 | 6.10 | 78 | 6.2
58Ni28+ | 2.07 | 110 | 6.25 | 96 | 6.7
18O8+ | 2.25 | 130 | 6.85 | 105 | (7.0)
40Ar18+ | 2.22 | 125 | (6.77) | 100 | (6.9)
238U91+ | 2.61 | 175 | (7.95) | 150 | (8.1)
86Kr33+ | 2.61 | 160 | 7.85 | 150 | (8.1)
abgebremste Ionen gemessen
(extrapolierte Amplitude)

## Slide 14: Commissioning alternative I (reduced scope)
1) Can we have “offline” D+ @ 4 MeV/u?
rapid commissioning without GSI beamtime
added bonus: available for other purposes, e.g. machine studies at ESR or HEST
… but cyclotrons or IH setups are costly and have no direct scientific benefit for GSI
http://cyclomed.tech
Development of a superconducting magnet for a compact Cyclotron for radioisotope production.
Luis García-Tabarés, Pablo Abramian, Jesús Calero, José L. Gutiérrez…
2) Can we reduce the scope of HITRAP?
HITRAP decelerates to 6 keV/u without cooling
EBIT and S-EBIT deliver HCI offline
still needs commissioning
106 cooled U92+ not possible, but…
… low intensity experiments possible
… HCI cooling for full scope can still be developed
Experiments
EBIT or S-EBIT