## Slide 1
Introduction to Spectrum Analysis Aufbau eines Spektrum-Analysators

## Slide 2: Contents
Introduction
RF measurements (time and frequency domain)
Signals of bunched beam current
The super-heterodyne concept and its application(s)
Spectrum Analyzer
Build Your Own Spectrum Analyzer

## Slide 3
Ibeam
Time Domain  Frequency Domain
Time domain: Recording of a voltage as a function of time:
Frequency domain: Displaying of voltage as a function of frequency:
Instrument:
Spectrum Analyzer
Fourier Transformation:
of time domain data
Instrument: Oscilloscope
Care: Fourier transf. contains amplitude and phase
courtesy company Keysight
voltage
time
frequency

## Slide 4: RF Measurements (Time Domain)
Oscilloscope: to observe signals in time domain
  periodic signals
  burst and transient signals with arbitrary waveforms
  application: direct observation of signals from a beam pick-up, from a test generator, or from other sources
  visualizes the shape of a waveform, etc.
  limited performance for the evaluation of non-linear effects.
Cathode Ray Tube (CRT) Oscilloscope
http://semesters.in/basic-principle-of-the-oscilloscope-notes-for-electronics-engineering-1st-year/

## Slide 5: Digital Storage Oscilloscope (DSO)
Signal processing based on fast ADCs and DACs
  Similar readout to an analog oscilloscopes, but better performance
    8…12-bit multi-GS/s ADCs, still, be aware of aliasing effects!
    Fast sampling oscilloscope require sufficient memory resources.
Typical time domain (TD) test setup: signal generator & digital oscilloscope:
  Device under test (DUT) characterization and trouble shooting
    Gaussian impulse, step, or arbitrary waveform (e.g. beam signal) as stimulus signal
    High impedance probe for measurements on the printed circuit board (PCB)
Device Under Test (DUT)
…and digital signal generator (AWG: arbitrary waveform generator)
50 GS/s, 10-bit AWG (Tektronix)
100 GHz bandwidth 240 GS/s oscilloscope (LeCroy)

## Slide 6: RF Measurements Methods
RF power detection: Cristal Detector (Schottky diode)
  Supplies a rectified (video) output signal proportional to the RF signal power
  Very broad frequency range few MHz to many GHz
  up to 90 dB dynamic range.
  Delivers no frequency nor phase information
Spectrum analyzer: to observe signals in frequency domain
  Sweeps in equidistant steps through a given frequency range
  application: observation of spectrum from the beam, or from a signal generator or RF source, or the spectrum emitted from an antenna to locate EMI issues in the accelerator tunnel, etc.
  Requires periodic signals
  Large dynamic range!
Vector signal analyzer (VSA), or FFT analyzer
  Acquires the signal, often after down-conversion, in time domain by fast sampling
  Further numerical treatment in digital signal processors (DSPs)
  Spectrum calculated using Fast Fourier Transform (FFT)
  Combines features of an oscilloscope and a spectrum analyzer: Signals can be observed directly in time or in frequency domain
  Contrary to the SA, also the spectrum of non-periodic signals and transients can be measured

## Slide 7: RF Measurements Methods
Vector Network Analyzer (VNA)
  Combines the functions of a vector spectrum analyzer (FFT analyzer), a RF sweep generator, and a S-parameter test set (directional coupler)
  Excites a Device Under Test (DUT, e.g. circuit, antenna, amplifier, etc.) network at a given Continuous Wave (CW) frequency, and measures the response in magnitude and phase => determines the S-parameters
  Covers a selectable frequency range by measuring step-by-step at subsequent frequency points (similar to the spectrum analyzer)
  Applications: characterization of passive and active RF components, Time Domain Reflectometry (TDR) by Fourier transformation of the reflection or transmission response, etc.
  The VNA is the most versatile and comprehensive tool in the RF laboratory!

## Slide 8
A (too) Simple Radio Receiver
Is this how does a ”traditional” analog radio works?
  It is, difficult to make precisely tunable narrowband, band-pass filters for high frequencies (~100 MHz)!!
  high frequency low-noise amplifiers are expensive!
  high frequency demodulators are not trivial.
    => direct detection of radio or RF signals is challenging!
broadband low-noise RF amp e.g. 87-108 MHz
tunable narrowband band-pass filter
RF amp
demodulator
audio amp

## Slide 9: The Super-Het Receiver
Introduce a non-linear element: the (frequency) mixer!
  ”down-convert” the RF band to a fixed ”intermediate” frequency (IF)
  requires a tunable local oscillator (LO)
  well manageable IF section:
    narrowband band-pass filter(s) (BPF) and amplifier(s)
  standards solution used in RF telecommunication
  Often multiple mixing stages are used in modern RF instruments (spectrum analyzer, vector network analyzer)
broadband low-noise RF amp e.g. 87-108 MHz
IF narrowband band-pass filter e.g. 10.7 MHz ± 90 kHz
IF amp
demodulator, e.g. FM PLL or AM diode detector
audio amp
tunable local oscillator (LO) e.g. 97.7-118.7 MHz
frequency mixer

## Slide 10
The RF mixer
Examples of different mixer configurations
A typical coaxial mixer (SMA connector)

## Slide 11
Mixer response in time domain
Input signals here:
LO = 10 MHz
RF = 8 MHz
frequency mixer
IF narrowband band-pass filter
Chart: Time (nanoseconds)
Chart: Time (nanoseconds)
Chart: Time (nanoseconds)
Chart: 
For more info, go to:
http://www.microwaves101.com/encyclopedia/mixerwaveforms.cfm

## Slide 12
Mixer response in time and frequency domain
Input signals here:
LO = 10 MHz
RF = 8 MHz
Mixing products at 2 and 18 MHz and higher order terms at higher frequencies
frequency mixer
IF narrowband band-pass filter
Chart: Time (nanoseconds)
Chart: Time (nanoseconds)

## Slide 13
The RF Mixer as Downconverter
Ideal mixer:
Frequency conversion
  fRF ≠ fLO: heterodyne receiver
  fRF = fLO: homodyne, demodulator
Real-world mixer:
  Image frequency:
RF
LO
IF
flo
fRF
fLO+fRF
fLO-fRF
f
A
fIM
upper sideband
lower sideband

## Slide 14
AM & FM Modulation
Beam current signals: transient (pulse-like) signals
RF signals: Sinusoidal signals (CW: continuous wave)
  High frequency carrier modulated with low frequency information
  In ring accelerators:
FM: Synchrotron oscillations
AM: Betatron oscillations

## Slide 15: The Super-Het Receiver
Introduce a non-linear element: the (frequency) mixer!
  ”down-convert” the RF band to a fixed ”intermediate” frequency (IF)
  requires a tunable local oscillator (LO)
  well manageable IF section:
    narrowband band-pass filter(s) (BPF) and amplifier(s)
  standards solution used in RF telecommunication
  Often multiple mixing stages are used in modern RF instruments (spectrum analyzer, vector network analyzer)
broadband low-noise RF amp e.g. 87-108 MHz
IF narrowband band-pass filter e.g. 10.7 MHz ± 90 kHz
IF amp
demodulator, e.g. FM PLL or AM diode detector
audio amp
tunable local oscillator (LO) e.g. 97.7-118.7 MHz
frequency mixer

## Slide 16: Simplified Spectrum Analyzer
based on the super-heterodyne principle
Switchable BW of the IF filter and video BPF (analog or digital) allows to improve the signal-to noise (S/N)-ratio

## Slide 17
The IF filter has to be a band-pass filter with a narrow frequency band
On the spectrum analyzer bandwidth are tunable
IF filter (BPF) tested with a VNA:
  Cavity filter
  central frequency: 101.8MHz
  band-width 3MHz
The Band Pass (IF) Filter

## Slide 18: Simplified Spectrum Analyzer
based on the super-heterodyne principle
Switchable BW of the IF filter and video BPF (analog or digital) allows to improve the signal-to noise (S/N)-ratio

## Slide 19: The RF diode
We are not discussing the generation of RF signals here, just the detection
Basic tool: fast RF* diode
(= Schottky diode)
In general, Schottky diodes are
fast but still have a voltage
dependent junction capacity
(metal – semi-conductor junction)
Equivalent circuit:
A typical RF detector diode
Try to guess from the type of the connector which side is the RF input and which is the output
Video output

## Slide 20
-20 dBm = 0.01 mW
Linear Region
This diagram depicts the so called square-law region where the output voltage (VVideo) is proportional to the input power
The transition between the linear region and the square-law region is typically between -10 and -20 dBm RF power (see diagram).
Since the input power is proportional to the square of the input voltage (VRF2) and the output signal is proportional to the input power, this region is called square- law region.
In other words:
VVideo ~ VRF2
The RF diode

## Slide 21: Modern Spectrum Analyzer in SA mode
based on the super-heterodyne principle
Today, the IF, demodulation, video and display sections of a spectrum analyzer are realized digitally
  Requires an analog-digital converter (ADC) with sufficient dynamic range
Switchable BW of the IF filter and video BPF (analog or digital) allows to improve the signal-to noise (S/N)-ratio

## Slide 23: Simplified Spectrum Analyzer
Receiving a 150MHz RF-signal from Function Generator
Local oscillator signal
  Frequency sweep from 100kHz to 100MHz in 1s
Output display
  0.1s/div. time base i.e. time corresponing to the frequency sweep
100kHz
100MHz
3MHz
Local oscillator signal
Output signal

## Slide 24: Further Reading
https://www.rohde-schwarz.com/de/applikationen/db-oder-nicht-db-educational-note_230850-15534.html
http://www.microwaves101.com/encyclopedia/mixerwaveforms.cfm
https://cds.cern.ch/record/1982420
https://cas.web.cern.ch/schools/tuusula-2018
https://cas.web.cern.ch/sites/cas.web.cern.ch/files/lectures/tuusula-2018/rfmeasurementstechniques.pdf
https://sites.google.com/view/uspas-dembd/home
https://indico.cern.ch/event/780638/timetable/#20190610

## Slide 25
Thank you!
Notes: 25

## Slide 26
Backup

## Slide 27: “dB or not dB”, that it the question
dezi-Bel: 1 dB = 0.1 B (Bel)
  Logarithmic scaling to compare large, e.g. power ratios:
  or large ratios of other quantities, e.g.:
dB ratio | P1/P2 | V1/V2
n x 10 dB | 10n | 10n/2
40 dB | 10000 | 100
20 dB | 100 | 10
10 dB | 10 | ~3.16
6 dB | ~4 | ~2
3 dB | ~2 | ~1.41
0 dB | 1 | 1
-3 dB | ~0.5 | ~0.71
-20 dB | 0.01 | 0.1
The 3 dB ratio (half power) is a common specification for the bandwidth
https://www.rohde-schwarz.com/de/applikationen/db-oder-nicht-db-educational-note_230850-15534.html

## Slide 28: “dBm” vs. “dB”
dBm is a logarithmic power unit
  based on dB and Pref = 1 mW
dBm can also be used as logarithmic voltage unit
  e.g. for Z0 = 50 Ω: Vref = 0.2236 V
dBm | P | V (RMS)
30 dBm | 1 W | 7.07 V
20 dBm | 100 mW | 2.24 V
10 dBm | 10 mW | 707 mV
6 dBm | 4.0 mW | 446 mV
0 dBm | 1.0 mW | 224 mV
-20 dBm | 10 μW | 22.4 mV
-60 dBm | 1.0 nW | 224 μV
-120 dBm | 1.0 fW | 224 nV
- 174 dBm | 4.0e-21 W | 0.446 nV
Noise power in a bandwidth B = 1 Hz at room temperature