## Slide 1: Motion prediction for online delivered dose verification in carbon ion therapy
K. Abu Mustafa, C. Galeone, M. Durante, C. Graeff, and L. Volz
Biophysics

## Slide 2: Moving tumors
M. Söhn, LMU
High spatial precision of particle therapy requires high precision image guidance

## Slide 3: Motion management in particle therapy
Active scanned beam
Target motion
ACRF Image X Institute
B. Marchand et al., EPAC 2000
Interplay effect
beam range
dose
lung
overshoot
Range changes
Motion mitigation inherently a 4D problem
Lateral tracking alone is not enough

## Slide 4: Treatment planning
Plan based on available motion information
T. Steinsberger, Doktorarbeit, GSI

## Slide 5: Irregular motion

## Slide 6: 4D RIDOS: a tool for real-time online dose calculation
C. Galeone et al. (2024) Phys. Med. Biol.

## Slide 7: Need for motion prediction
Motion inference from optical tracking system (OTS)
Lag of the system
For RIDOS: ms speed is important to ensure real-time dose assessment
Modified from presentation by M. Desplanques (CNAO)
Design LSTM for tumor motion prediction

## Slide 8: Model

## Slide 9: Data
Experiments at CNAO (Italian carbon ion therapy center)
Carbon plans irradiated to moving IC detector array
Recording of motion traces with optical tracking system
Sine motion with irregular variations

## Slide 10: Data
9 recorded motion traces, 7/2 train test split
Total of ~1h of motion signals

## Slide 11: Data augmentation
Data augmentation (amplitude variation, baseline drifts) for more robust model

## Slide 12: Results: Motion traces (test 1)
Prediction: 8s input, 250ms prediction (equivalent to 10 data points) into the future
Maximum error: 2.44 mm
MAE: 0.51 mm

## Slide 13: Results: Motion traces (test 2)
Prediction: 8s input, 250ms prediction (equivalent to 10 data points) into the future
Maximum error: 9.33 mm
MAE: 0.45 mm

## Slide 14: Results: comparison of dose calculation
Real motion trace
Gamma index: 98.8%
Predicted motion trace
Gamma index: 98.8%
No significant difference between predictions based on real and predicted motion
Use predicted motion trace as input to 4D-RIDOS

## Slide 15: Conclusion
Motion is a challenge for precise ion beam treatments
Online dose verification is needed to ensure correct dose delivery
We designed an LSTM model to predict motion from optical tracking system to be used as input for real-time delivered dose verification
So far: promising results on test data (still idealized)
Next:
  Application to real breathing curves
  Anatomically informed model design