## Slide 1
Volker Tympel, HI-Jena
Sept. 17th, 2025
SPARC 2025
Topical Workshop
Cryogenic Current Comparators:
Status and challenges

## Slide 2
CCC?
charged
particle
beam
B
Principle of the
Cryogenic
Current
Comparator
Low intensity beam
B=0
-i
i
Thermal noise reduction
Superconductivity
  DC transformer (=> un-bunched beam)
  Magnetic shielding
  Superconducting Quantum Interference Device (SQUID)
=> absolute values
Calibration certificate
current source
=> non-destructive

## Slide 3
CCC?
Low intensity beam
Screen shot
CCC running @
CERN-AD.
CCC is activated in the control room.
2017
Comparison test CCC against
Parametric Current Transformer and
Ionization Profile Monitor @
CRYRING@ESR in 2021

## Slide 4
CCCs
Evolution until 2024
GSI
1996
CERN-
Nb-CCC
2015
BESSY
2009
FAIR-
Nb-CCC-xD
2017
Pb-Coreless CCCs V1…V8, EMI-problems
unstable in the field of application
Pb-DualCoreCCC
Pb-DCCC-xD
2024
Pb-DCCC-Sm-200
2022
Pb-DCCC-Sm-300
2023
Pb-CCC-xD-V8-2
2024
BMBF project:
05P15SJRBA
BMBF project:
05P18SJRB1
BMBF project:
05P21SJRB1
PhD theses
R. Geithner
2013
PhD theses
M. Stapelfeld
2022
PhD theses
M. Fernandes
2017
PhD theses
F. Kurian
2016
PhD theses
D. Haider
2022
PhD theses
L. Crescimbeni
in progress
IBIC‘98
Farady Cup Award
to A. Peters (GSI)
Noise: 250 pA /√Hz
BLD 100 mm
BLD 150 mm
BLD 40 mm
Tested @ GSI
Finally

## Slide 5
Status: Lab Limits
Sm version
Current resolution
White noise
Spectral
current noise
density
pArms/√Hz
Temperature
4.2 K → 1.9 K
Noise reduction:
0.67
Real part of the
inverse impedance
fixed for a
core material
fixed by
SQUID
Inductance
100 µH → 200 µH
Noise reduction:
≈0.7
Fluctuation–dissipation theorem*
* R. Geithner, PhD theses, Jena 2013
Confirmed Tympel, et al., High Inductance Cryogenic Current Comparators for Beamlines
IEEE Trans. Appl. Supercon. Vol. 34, Is.: 3, May 2024

## Slide 6
Status: Lab limits
Current resolution
White noise: 1 pArms/√Hz @ 300 µH, 2.1 K
Pulse / ramp
Pulse: 200 µs (Sm-200 un-damped)
Ramp: 10 ms (Sm-300 un-damped)
0.5 nApp
Sm version

## Slide 7
Status: Lab limits
Current resolution
White noise: 1 pArms/√Hz
Pulse / ramp: 500 pApp
Frequency bandwidth
from DC (to be defined!)
Low Frequency Noise
Measured @ 4.2 K
LC meander resonance
20 dB per decade
=> like Brownian noise
1/f2
Current or voltage in dB
=> 20 ·lg10(xout / xin)
White noise
Sm version

## Slide 8
Status: Lab limits
Current resolution
White noise: 1 pArms/√Hz
Pulse / ramp: 500 pApp
Frequency bandwidth
from DC
to 1.5 MHz
Pb-DCCC-Sm-300
final damped version
Calibration U=>I
via factor
via deconvolution
CERN-Nb-CCC
FAIR-Nb-CCC-xD
Pb-DCCC-xD
Deconvolution:
Digital data processing necessary !
Low sensitive SQUID
Sm version

## Slide 9
Status: Lab limits
Current resolution
White noise: 1 pArms/√Hz
Pulse / ramp: 500 pApp
Frequency bandwidth
from DC
to 1.5 MHz
Pb-DCCC-Sm-300
final damped version
Dynamics
Low sensitive SQUID
7x slew rate extension
3 SQUIDs => 3 zero lines:
Digital data processing necessary !
Sm version

## Slide 10
Status: Lab limits
Current resolution
White noise: 1 pArms/√Hz
Pulse / ramp: 500 pApp
Frequency bandwidth
from DC
to 1.5 MHz
Dynamics
7x slew rate extension
Nb-CCCs
10-5
Pb-foil CCCs
10-8
Magnetic shielding
1000x
Sm version

## Slide 11
Status: Systems
Cryostats
CERN-AD
since 2017
FAIR #1
since 2021,
now alteration:
125 => 150 mm
FAIR #2
order placed
CRYRING@ESR
ongoing tender
Beamline Ø: 150 mm
Jena-GFK
since 2024
without He Reliquefier
CCC sensor versions:
-xD
100 mm
40 mm
-Sm
-mD

## Slide 12
Challenge: FAIR
Planning for FAIR
xD version
FAIR #1 / #2

## Slide 13
He Reliquefier
PT425-RM cold head
Rotary valve
He 725 bar
@ 1.4 Hz
He 7 bar
He 25 bar
150 mV
40 mV
10 mV
Challenge: Microphonics

## Slide 14
He Reliquefier
?
Challenge: Microphonics

## Slide 15
Digital data processing
Channel Red SQUID (LP-filtered)
Raw data
24 bit delta-sigma ADC
10 kSample/s
Calibration: 0.24 nA/mV
85 nA
Challenge: Microphonics

## Slide 16
5 min
Data sample
6.7 mHz
Challenge: Microphonics
Digital data processing
+

## Slide 17
PT425-RM cold head
PT425-RM cold head
PT425-RM cold head
PT425-RM cold head
PT425-RM cold head
Vacuum pumps ?
Air conditioner ?
?
?
?
?
?
?
Digital data processing
40 mV
10 mV
Challenge: Microphonics

## Slide 18
Electrical coupling
Asynchronous motors
(Induction motor)
Mechanical coupling
Rotor
=> Slip frequency
1 % slip
0.6 % slip
Stator
50 Hz
10 mV
Challenge: Microphonics
Digital data processing

## Slide 19
EAH-Jena University of Applied Sciences
BMFTR-Project:
05P24SHA 2024-27
Department of Electrical Engineering and Information Technology
Prof. Dr.-Ing.
Burkart Voss
M. Eng.
Svenja Kolbe
Challenge: Microphonics
Digital data processing

## Slide 20
ErUM Matter call in 2024
special ½ meter Compact CCC system
BMFTR-Project: C4CRY
05K25SJA 2025-28
Challenge: ½ meter
CRYRING@ESR
Open PhD position!
Contract partner: University of Jena
Place of work: GSI Darmstadt

## Slide 21
Present situation with
DC beams:
Diagnostics relies only
on beam intercepting devices.
New Pb-DCCC for slow extraction from SPS to
North Area (TT-20). (J. Tan @ IBIC 2024)
Why a CCC?
DC & AC,
<10 µA,
abs. current,
non-destructive
Sensor delivery: 2026 ?
Challenge: CERN
2nd CERN system
2nd
1st

## Slide 22
Summary
Status: Clear Design & Lab Limits
Challenge: Data Processing (Microphonics, 3 SQUIDs)
Complete Systems
Growing: Cooling Costs (Beamline Cryostat, He Reliquefier)
Thank you for your attention!