## Slide 1
Rohit Tyagi, J. Viehmann, A. Krone, A. Hans, A. Ehresmann
University of Kassel
23-01-2026
Current status of VUV-EUV VLS-grating photon spectrometer at CRYRing

## Slide 2
A first test experiment
Reiß et al., Phys. Scr. T166, 014031 (2015)
impact on -gas target
Spectral range of 340 nm – 550 nm
Non UHV-compatible 1m-Seya-Namioka spectrometer
Observation through window
Spectrometer not in focus to source volume
Obstacles due to geometric constraints and requested vacuum conditions
Result:
In spite of obstacles signal-to-noise ratio good  Worthwhile to build a EUV/VUV spectrometer
2

## Slide 3: Why VLS ?
3
NIS (Normal incidence spectrometer) system the spectra must be recorded and appended at 7 different grating positions
VLS grating Full spectra can be taken in single shot
Ar and fluorescence emission 2 keV electron impact
Aberration-corrected (defocus, coma, astigmatism) by design
Flat-field imaging allows 2D detectors enabling imaging spectrographs
More complex and costly fabrication of VLS gratings
Optimize for wavelength band requiring multiple gratings for wide coverage.

## Slide 4
Ion-Atom Collision
/
/
/
/
Excitation
Pure Ionization
Transfer Processes
Radiative
relaxation
UV/VUV
Spectroscopy

## Slide 5
Motivation: (Single or multiple) electron transfer/ capture to/ by highly charged ions
2s
2p
Ne7+
EI+
3s
Ne
4p
2p
2s
Projectie
Target
2s
2p
Ne6+
EI+
3s
Ne+
4p
2p
2s
Projectile fluorescence (for NeVII in the range of 46.5 – 97.3 nm, for higher excited states above 198 nm)
Possibility (a): Surplus energy released by Capture fluorescence
2s
2p
Ne7+
EI+
3s
Ne
4p
2p
2s
Projectile
Target
2s
2p
Ne6+
EI+
Target fluorescence
Possibility (b): Surplus energy released by Correlative Excitation in target
3s
Ne+*
4p
2p
2s
Measurable?

## Slide 6
www.arxiv.org/pdf/2511.20537 Nov 2025
Precision experiment for QED
Cont..
The e-e interaction enables this capture channel to occur and another
valence electron will be excited due to the recoil force
PHYSICAL REVIEW A102,042820(2020)

## Slide 7
UV/VUV detection @ CryRing
10-120 nm, 100-300 nm
Remotely controllable adjustment stage
Slit unit, if used in focus of the variable line spacing grating (35.8cm away from the grating center)
Position-sensitive photon detection unit no. 1, here EUV-detection unit
Position-sensitive photon detection unit no. 2, here VUV-detection unit
Pumping unit

## Slide 8
Different use cases:
Grating (1) target fluorescence and grating (2) projectile fluorescence downstream of the target
Both grating-detector combinations are used together downstream of target area for projectile fluorescence
Alternate use of the grating-detector combinations by stage shift along ion beam for target fluorescence
Grating and Detector positions and characteristics
EUV: 300 l/mm (avg) Al + MgF2-coating focal length 35.8 cm 50 x 30 mm active area
VUV: 1300 l/mm (avg) Au coating focal length 35.8 cm 50 x 30 mm active area

## Slide 9
Installation
Preliminary test
Integration of spectrometer control and data acquisition systems into the GSI infra structure. Iteration to solve technical limitation
User operation
Beam time (Optimistic :))
Work plan including milestones and status

## Slide 10
Facts and challenges from beamtime
background
Krypton
Xenon
Spectra measured without slit unit in about 45 min
Target: Kr/ Xe at 25bar, nozzle 25 microns, trumpet-shaped; projectile: Ne7+ @ 5MeV/u
Clear target-specific spectra detected  signal to noise even at this unfortunate spectrometer position is o.k.
10

## Slide 11: What are we looking for?
Kr (4p65s0)
S Tsurubuchi et al 2003 J. Phys. B: At. Mol. Opt. Phys. 36 2629
5s4p5  4p6
(116.4 nm)
(123.5 nm)

## Slide 12
Kr (4p65s0)
S Tsurubuchi et al 2003 J. Phys. B: At. Mol. Opt. Phys. 36 2629
Target: Kr/ Xe @ ~25bar, nozzle ~25microns, trumpet-shaped
Projectile: Ne7+ @ 5MeV/ion
Experimental run @ GSI
Cont..

## Slide 13: Kr ionic species lines
32.7
1.93
0.87
0.52
1.15
0.91
1.5
Ne
Kr
Kr I
(Hartree unit)
Roothaan-Hartree-Fock Atomic wave function by Clementi and Roetti
,
147.0 nm
129.6 nm
Xe I: 5p5(2P3/2)6s 2[3/2]1  5p6 1S0
Xe I: 5p5(2P1/2)6s 2[1/2]1  5p6 1S0

## Slide 14: Previous set-up New set-up
Dectector
Grating
slit
Interaction
region
Dectector
Grating
slit
Interaction
region
400-450mm
50mm
Problem | Possible Cause
Very low signal | Distance between the grating and interaction region is large
Poor resolution | Slit needed to be removed, not in grating focus
Calibration | Spectra used might have slightly different conditions
Changes | Possible improvement
The distance between the grating and interaction region | More count rate
Use of slit or focused on interaction region | Might improve resolution
More gas density | More count rate
Projectile energy seems high for capture process
We should try for different projectile energy
Diagnostics and upgrade

## Slide 15
Projectile energy 5 MeV per u
Projectile energy 1 MeV per u
Kr Target density : 6.46 xatoms per
Projectile Ne7+ charge state distribution after collision
LISEcute++

## Slide 16
A. Hans C. Küstner-Wetekam L. Marder
A. Krone D. Bloss C. Honisch N. Kiefer J. Viehmann …
Institute of physics U Kassel
First detection of Kr- and Xe-target fluorescence after slow Ne7+ ion collision
Hints to fluorescence from target ions
There is room for improvement, but exciting, science seems very well feasible
We hope to obtain good signal to noise ratio in new configuration
Summary
Current status and future plan for VLS spectrometer
Great support/ collaboration by
A. Gumberidze S. Hagmann E. O. Hanu
M. Lestinsky N. Petridis E. Menz U. Spillmann
Th. Stöhlker …
GSI
P.-M. Hillenbrand
C. Brandau
S. Schippers
…
JLU Giessen
Monetary support by
is gratefully acknowledged
धन्यवाद || Thank You
Danke 

## Slide 17
131 mm
358 mm
55 mm
100 mm
Special bellow: Total relax length not more than 70 mm
Idea: DN63 bellow both end with DN 100 Flange.
Grating:

## Slide 18
2s
2p
EI+
3s
4p
2p
2s
2s
2p
EI+
3s
4p
2p
2s