## Slide 1
Bunch Struktur Messungen am UNILAC --- Status Report ---
Operator-School, 30.10.2025
Beam Instrumentation Abteilung:
Peter Forck, Stephan Klaproth, Timo Milosic, Thomas Sieber, Rahul Singh
LINAC Abteilung:
Maksym Miski-Oglu, Uwe Scheeler, Nimue Schmidt,
Winfried Barth, Simon Lauber et al.

## Slide 2
Messungen longitudinal Parameter:
Beispiele zur Simulation des longitudinaler Phasenraums
Generelles zur Bunch-Länger Messung am UNILAC
Monitor 1: Bunch Shape Monitor (BSM) von russischen Institut INR
Monitor 2: Fast Faraday Cup (FFC); Design vom amerikanischen Fermilab
Erste Ansätze zur Bestimmung des longitudinalen Emittanz
Vergleich von BSM und FFC
Zusammenfassung und Ausblick
Gliederung

## Slide 3
Bunch shapes are very non-Gaussian
Reason: Acceleration on maximum of rf-wave to enable maximum energy gain
by KONUS beam dynamics at CH-cavities (KoNuS = Kombinierte Null Grad Struktur)
Bunch Shape by KONUS Beam Dynamics
Example
325 MHz Proton LINAC:
Longitudinal phase space
simulation:
University Frankfurt
Deformed phase space
Complex bunch shape
Sensitive to small variation
of cavity setting (U and )
Emittance minimization only
by correct setting
MeV  =37%
MeV  =8%
E=MeV  =12%
Phase :
span 900
= 0.77 ns
Phase :
span 400
= 0.34 ns
Phase :
span 200
= 0.17 ns

## Slide 4
Bunch Shape by KONUS Beam Dynamics
Simulation example HSI exit:
High current U4+, 1.4 MeV/u
Longitudinal phase space
Measurement example HLI at UN6:
Low current, 1.4 MeV/u
Measurement: Tomographic reconstruction
W. Barth, S. Yaramishev et al. LINAC’06
Phase  [degree]
0
9
-9
-18
0
18
0.45
Energy spread E/E [%]
-0.45
0.9
-0.9
S. Lauber et al., IPAC’19 & Phys. Rev. Accel. Beams, 23, 114201 (2020)
Phase  span 360 = 2.8 ns @36 MHz
Bunch shapes are very non-Gaussian
Reason: Acceleration on maximum of rf-wave to enable maximum energy gain
by KONUS beam dynamics at IH-cavities (KoNuS = Kombinierte Null Grad Struktur)
Phase  span 720 = 1.85 ns
span 0.8 %

## Slide 5
Example: 1010 U28+ at 11.4 MeV/u
injection plateau 150 ms, η = 0.94
Longitudinal Schottky at harmonics h = 117
Momentum spread variation:
p/p  (0.6... 2.5) 10-3 (1)
Momentum spread  p/p0 after multi-turn injection
& de-bunching of t < 1ms duration
to stay within momentum acceptance during acceleration
Method: Variation of buncher voltage
i.e. rotation in longitudinal phase space
minimizing of momentum spread  p/p0
Longitudinal Schottky for Momentum Spread p/p0 Analysis
time or phase
 = p/ p
long. phase space
LINAC bunches at injection:
from LINAC
35 kHz
8 dB
p/p  2.5 10-3
p/p  1.3 10-3
p/p  0.6 10-3
bunched
only drift
de-bunched
6 dB
6 dB
fcenter = 24.9 MHz
fh  hp
synchrotron
LINAC
buncher
time or phase
 = p/ p
long. phase space
de-bunching
De-bunching after some ms:
Schottky
injection
extraction
Schottky
pickup

Usum
Uleft
Uright
However:
Schottky show only the
stable stored beam
not UNILAC beam p/p

## Slide 6
Bunch Structure at low Ekin: Not possible with Pick-Ups
Pick-ups are used for:
precise for bunch-center relative to rf
But:
For  << 1  long. -field significantly modified:
Example: Comparison pick-up – Bunch Shape Monitor:
HLI: Ar beam, 1.4 MeV/u ( = 5.5%) , frf = 108 MHz
 the pick-up signal is insensitive
to bunch ’fine-structure’
ampl.
R
  R / 
1/ frf = 9.2 ns
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM
 tbunch=1.3 ns

## Slide 7
Messungen longitudinal Parameter:
Beispiele zur Simulation des longitudinaler Phasenraums
Generelles zur Bunch-Länger Messung am UNILAC
Monitor 1: Bunch Shape Monitor (BSM) von russischen Institut INR
Monitor 2: Fast Faraday Cup (FFC); Design vom amerikanischen Fermilab
Erste Ansätze zur Bestimmung des longitudinalen Emittanz
Vergleich von BSM und FFC
Zusammenfassung und Ausblick
Gliederung

## Slide 8
Bunch Structure using secondary Electrons for low EkinIons
Secondary e− liberated from a wire carrying the time information.
→ Bunch Shape Monitor (BSM)
Working principle:
Insertion of a 0.1 mm wire at -10 kV
Emission of secondary e− within ps
 accelerated toward an rf-deflector
rf-deflector as ’time-to-space’ converter
Detector behind a thin slit
Scanning method:
Shift of the phase  for each macro-pulse
Resolution  10 ps  1o @ 280 MHz
SEM: Secondary Electron Multiplier

## Slide 9
Realization of Bunch Shape Monitor at GSI
UNILAC BSM
resonator housing
lens HV /steerer
target HV
actuator
rf-coupling
viewing ports
BEAM
µ metal
~ 1m
frequency tuner
rf-pickup
magnets
e- multiplier
Remark:
Insertion length in beam direction: 129 mm
Alignment with flange accuracy sufficient
Installation at UCW
HELIAC (cwLINAC)
Installation of the three BSMs:
Presently: UN6 (HLI), X2, UCW (after cavities)
Planned: UT2, TK7, UCW (after cavities)
Installation at UCW:
Electronics
CF63

## Slide 10
Performant GUI:
Control of various monitor parameters
Bunch shape averaged over macro-pulse
Time resolved bunch with 1 µs steps
Bunch Shape Monitor: Graphical User Interface
Phase  span 1800 = 4.6 ns
Remarks:
BSM system responsible:
Thomas Sieber, Pere Forck (deputy)
Control parameters are fixed during measurements and are aligned by BEA members
Moreover, variation during macro-pulse  should be observed and interpreted !
Time [µs]
Phase [deg]
Bunch shape (entire macro-pulse)
100@108MHz
=157 ps
BSM signal of
macro-pulse
Scan range
Control parameters
Control parameters
Transformer
signal of
macro-pulse

## Slide 11
Messungen longitudinal Parameter:
Beispiele zur Simulation des longitudinaler Phasenraums
Generelles zur Bunch-Länger Messung am UNILAC
Monitor 1: Bunch Shape Monitor (BSM) von russischen Institut INR
Monitor 2: Fast Faraday Cup (FFC); Design vom amerikanischen Fermilab
Erste Ansätze zur Bestimmung des longitudinalen Emittanz
Vergleich von BSM und FFC
Zusammenfassung und Ausblick
Gliederung

## Slide 12
Fast Faraday Cup FFC for Bunch Shape Measurement
Design at Fermilab (USA) by D. Sun & A. Shemyakin,
see J-P Camiero et al. Int. J. of Mod. Phys. A, vol. 34, no. 36, 2019
Fast Faraday Cup for direct bunch charge measurement:
Round 50  co-axial line
Bandwidth  20 GHz
Collimator for beam with  0.8 mm hole
Collimator acts as a shield for -field
Secondary e- suppression by bias voltage
Mounting on stepping motor drive
Electronics:
Commercial Bias-Tee (bandwidth 10 GHz)
Broadband amplifier
Fast oscilloscope with at least 8 GHz bandwidth
Ubias
N-connector
0.8 hole for beam
N- connector
Time resolution:
61 ps for β = 5%
22 ps for β = 15%
Determined by:
Gap between inner&outer conductor
Electronics bandwidth

## Slide 13
FFC display with TIRESCO:
Readout of Tektronix scope MSO64B (6 GHz, 25 GS/s)
Bunch shape along macro pulse
Here: variation during macro-pulse
Average along macro pulse
Status:
Software by Timo Milosic
Almost finished
Fast Faraday Cup FFC: Graphical User Interface
Installation:
Presently:
UN6, UCW, US3, X2
Planned:
UN6, UCW, X2, US3, UT2, TK7
Broadband electronics and DAQ in preparation
Care: FFC might be destroyed by high beam power
Bunch shape along macro-pulse
Bunch shape
Integrated over macro-pulse
At time t1 & t2 during pulse
Phase  span 3600 = 9.26 ns

## Slide 14
HLI @ UN6: Ar8+, 1.4 MeV/u, 50 µA ( %)
Horizontal movement by 8 mm of BSM ( 0.1 mm wire) and FFC ( 0.8 mm hole)
 Horizontal – longitudinal correlation even straight after IH (unexpected)
 Further investigations required
Comparison FFC and BSM, Bunch Shape Variation
FFC
Phase  span 2000 = 5.14 ns
BSM
Phase  span 2000 = 5.14 ns
9 mm hori.
movement

## Slide 15
Messungen longitudinal Parameter:
Beispiele zur Simulation des longitudinaler Phasenraums
Generelles zur Bunch-Länger Messung am UNILAC
Monitor 1: Bunch Shape Monitor (BSM) von russischen Institut INR
Monitor 2: Fast Faraday Cup (FFC); Design vom amerikanischen Fermilab
Erste Ansätze zur Bestimmung des longitudinalen Emittanz
Vergleich von BSM und FFC
Zusammenfassung und Ausblick
Gliederung

## Slide 16
Novel investigations of longitudinal emittance:
Location with large dispersion
Bunch shape for different positions
Realized at US3 and X2
e.g. 30th July 25 at X2 using E10 as buncher
Idea for longitudinal Emittance Measurement
Calibration of dispersion 0.900.1 m using at X2 via slight acceleration and profile measurement
X2 July 2025: 0.900.1 m  mm for % (using )
N. Schmidt et al, IBIC’25 & Master Thesis

## Slide 17
Idea for longitudinal Emittance Measurement
Experimental confirmation  position scan of FFC (30th July 25):
Beam from HLI at X2: E= 11.4 MeV/u (), µA
Variation of E10 as buncher
Result:
Method is appropriate
Dispersion measurement by energy variation and SEM-grid detection
Strong deformation of longitudinal phase space
Evaluation is ongoing
Phase  span 3600 = 9.25 ns
about
using measured
0.900.1 m

## Slide 18
Long. Emittance by linear Transformation using a Buncher
Longitudinal focusing:
Variation of the bunch shape by a rf-buncher
 components 5 and 6 from 6-dim phase-space
Transversal correspondence:
Quadrupole variation
 Transfer matrix of buncher & drift:
with focal length
 Variation of buncher amplitude :
 Different bunch width at Bunch Shape Monitor
 Emittance determination by linear transformation with
assumption: Bunch has Gaussian form  phase space is given by an ellipse

## Slide 19
Result of a longitudinal Emittance Measurement
Example UNILAC-TK5: Voltage variation at E10 for 11.4 MeV/u Ni14+ beam, 31 m drift:
This example: The resolution is better
than 50 ps or 2o for 108 MHz
Typical bunch length at proton LINACs:
bunch  0.1 to 0.5 ns
Determination of longitudinal emittance possible
However: Bunch has non-Gaussian shape!
Application for synchrotron injection:
Shaping of longitudinal phase space by buncher
i.e. long bunches  low momentum spread
to match to the synchrotron long acceptance
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM
extraction
injection
p synchrotron
de-bunched
multi-turn
injection

## Slide 20
Longitudinal Emittance using tomographic Reconstruction
original object
Iterative projection
& back-projection
after sufficient
iterations
1st back-
projection
Filtered back projection:
Iterative process by redistributing
the 2-dim image and considering the
differences to the previous iteration step.
Tomography: Generation of a n-dim image
from many lower-dim projections
Medical application: 2-dim reconstruction using sufficient 1-dim projections

## Slide 21
Longitudinal Profile & Emittance: Tomographic Reconstruction
Example GSI HLI: Voltage variation at buncher for 1.4 MeV/u Ar9+ beam
Measurement examples with BSM
for different buncher settings using ‘Non-Negative Least Square’ solver including space charge
S. Lauber et al., Phys. Rev. Accel. Beams, 23, 114201 (2020)
Result:
Non-regular phase-space distributions are recovered.
Phase  span 2000 = 5.1 ns
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM
Measurement
at BSM
Phase  span 720 = 1.85 ns
span 0.8 %
Reconstruction at HLI output

## Slide 22
Messungen longitudinal Parameter:
Beispiele zur Simulation des longitudinaler Phasenraums
Generelles zur Bunch-Länger Messung am UNILAC
Monitor 1: Bunch Shape Monitor (BSM) von russischen Institut INR
Monitor 2: Fast Faraday Cup (FFC); Design vom amerikanischen Fermilab
Erste Ansätze zur Bestimmung des longitudinalen Emittanz
Vergleich von BSM und FFC
Zusammenfassung und Ausblick
Gliederung

## Slide 23
Resolution measurement at X2 (30th July 2025):
At X2 with beam from HLI:
11.4 MeV/u (), µA
Comparison FFC and BSM
Interpretation:
BSM & FFC can be used for measurement
BSM: - Better resolution (about  10 = 25 ps)
- Long measurement time due to scan
FFC: - Less resolution (about  50 = 125 ps)
depending on electronics & cable bandwidth
- Fast measurement within one macro-pulse
Applications:
FFC  Fast information during operation
BSM  Precise measurement
Goal: Tomographic reconstruction etc.
Phase  span 320 = 0.82 ns
Centre of beam short bunch
FFC
BSM
Phase  span 320 = 0.82 ns
Side of beam  long bunch
BSM
FFC

## Slide 24
Messung longitudinal Parameter ist wichtig:
Bisher sehr wenig longitudinale Messung am UNILAC durchgeführt
Komplexe KONUS Strahldynamik ist sehr sensitiv auf und Werte
Optimale Parameter (insbesondere ) für SIS-Injektion erfordert genauer Einstellung
Strahldynamisches Verständnis muss erarbeitet werden
Zusammenfassung und Ausblick
Zwei Messsysteme detailliert getestet (Fast Faraday Cup & Bunch Shape Monitor):
FFC: - Messung innerhalb eines Makro-Pulses
- Auflösung ca. 100 ps (abhängig von , Elektronik & Kabel)
BSM: - Scan-Verfahren, benötigt min.  100 Pulse, komplexerer Monitor-Aufbau
- Auflösung ca. 10 ps = 0.50@108MHz
- Geeignet für genaue Messung und Tomographie etc.
Einbau:
FFC: UN6, UCW, US3, UT2, TK7oder 8, X2
BSM: 3 Monitore vorhanden;
bisher: UN6, UCW, X2; zukünftig: UCW, UT2, TK7 oder 8 (parallel zu FFC)
Tests durch LINAC- und BEA-Team in 2026 bei Engineering Runs & Operation bei Bedarf
Bemerkung: Variationen über den Makro-Puls detektiert; dies ist auch bei den Phasensonden sichtbar
(passender GUI wird hoffentlich installiert)  bitte im Operating beobachten.
Vielen Dank für die Aufmerksamkeit. Gibt es Fragen & Bemerkungen?

## Slide 25
Back-up slides

## Slide 26
Locations and measurements
Installation at UT1 (pLinac BSM) and TK8 (UNILAC BSM) to investigate beam behind UE1 and TK debuncher (SIS18 injection optimization)
Longitudinal emittance studies by reproduction technique
Investigate effect of bunch shift during macropulse in X2
Investigate position dependent bunch shape (including FFC measurement)
BSM at GSI, current situation:
3 BSM at GSI: UNILAC- (2017, 108, 216, 325 MHz); cw- (2020, 108 MHz), pLinac- (2020, 325 MHz)
UNILAC and cw-BSM frequently (and currently) in operation at HLI/HELIAC
Recommendation 22/23 (MAC / Technical Integtration Team) to investigate UNILAC longitudinal beam properties (BSM +FFC)  meeting LIN + BEA  install two BSM in UNILAC+TK
pLinac BSM (after commissioning at HLI) modified to 216 MHz and tested in X2
BSM 1 (+FFC)
216.816 MHz
UT1
BSM 2 (+FFC)
108.408 MHz
TK8
UNILAC
Single gap
resonators
To M- caves
Notes: 26

## Slide 27
Low Velocity Effect: General Consideration
Lorentz transformation of single point-like charge:
Lorentz boost and transformation of time:
Trans. Elab.-frame of a point charge:
Long. E|| lab.-frame of a point charge:

## Slide 28
Fast Faraday Cup FFC Bandwidth
Design at Fermilab (USA) by D. Sun & A. Shemyakin,
see J-P Camiero et al.
Int. J. of Mod. Phys. A, vol. 34, no. 36, 2019
Fast Faraday Cup for direct bunch charge measurement:
Round 50  co-axial line
Bandwidth  20 GHz
Ubias

## Slide 29
SEM-grid described at: Juan Herranz et al., IBIC‘22
Measurement of Dispersion at X2
Experimental determination dispersion at X2 using :
Slight acceleration with E10
ToF measurement with Pasensonden in UT
Profile measurement at X2 with 0.5 mm wire spacing, 64 wires (from pLINAC)
Beam from HLI at X2: E= 11.4 MeV/u (), µA
Exp. errors determined by transmission variation
0.900.1 m

## Slide 30
Longitudinal Profile & Emittance: Tomographic Reconstruction
Example GSI LINAC: Voltage variation at buncher for 1.4 MeV/u Ar9+ beam
Arbitrary phase space shapes can be recovered by reconstruction techniques (within some uncertainty)
Phase space at buncher in discrete areas with m grid points with intensities  called mapping
With particle tracking codes transferred to BSM for different buncher setting
Projection to time axis for all buncher setting to form a vector
 Can be interpreted as vector-matrix equation
Matrix of dimension (n x m) = length of all phases x number of grid points m
resolution for phases : 100 , number of measurement:  n = 5000
Best fit by minimizing for the best phase space distribution
by Non-Negative Least Square
(NNLS) method
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM
S. Lauber et al., Phys. Rev. Accel. Beams, 23, 114201 (2020)
Phase 

## Slide 31
Performant GUI:
Control of various monitor parameters
Bunch shape averaged over macro-pulse
Time resolved bunch with 1 µs steps
Bunch Shape Monitor: Graphical Unser Interface
Bunch shape (entire macro-pulse)
50 @ 108MHz=129 ps
Signal of
macro-pulse
Scan range
Control parameters
Control parameters
Phase  span 600 = 1.54 ns
Remarks:
BSM system responsible:
Thomas Sieber, Pere Forck (deputy)
Control parameters are fixed during
measurements and are aligned by BEA members
Time [µs]
Phase [deg]

## Slide 32
Longitudinal Profile & Emittance: Algebraic Reconstruction
Example GSI HLI: Voltage variation at buncher for 1.4 MeV/u Ar9+ beam
Measurement examples with BSM
for different buncher settings:
S. Lauber et al., Phys. Rev. Accel. Beams, 23, 114201 (2020)
Measurement at BSM
Example
Nr. 1
Example
Nr. 2
Example
Nr. 3
Result: Non-regular phase-space distributions recovered
Test by changing Buncher 2  trustful reconstruction
Phase  span 2000 = 5.1 ns
LINAC
RFQ
source
Bunch Shape Monitor
B.1
Buncher
BSM 2
Measurement
at BSM 2
B.2