## Slide 1: Dosimetry with Uranium beams
E. Kozlova, A. Sokolov, S. Kundumattathil Mohanan, T. Radon, E. Mustafin

## Slide 2: CERN Radiation Monitoring Electronics (CROME) @ FAIR
Surveyed Area
Limit: 0.5 uSv/h
Controlled Area
Limit: 3 uSv/h
Dose Monitoring
Part of Personal Access System (PAS)
Measuring Station
PMI + CROME – were tested during Machine Studies in 2024 and approved themselfs to be the perfect candidates for dose measurements at FAIR facility

## Slide 3: Dosimeters and Electronics
CROME PC
WENDI II
FHT
192
FHZ 632L-H10
CROME R
CROME PC
FHT
190
FHT6020
WENDI II
FHT
192
FHZ 632L-H10
GSI‘s passive dosimeters
Thermo Fisher‘s active dosimeter

## Slide 4: Experimental Setup
990 MeV/u
´cc
Al-Target 10x10x21.1 cm³
Set of 114 plates
C
D
90°
40°
Target
U+73@ 990 MeV/u slow extracted beam. The transmission from SIS to Cave A was about 17%.
SEETRAM
WENDI II
FHT 192
FHT 190
FHZ 632
WENDI II
GSI-Sphere
1
2

## Slide 5: Dose rates/FLUKA Simulations
Total ambient dose
Neutron ambient dose

## Slide 6: Neutron Detectors
GSI‘s passive neutron dosimeter
Wendi-II active dosimeter
Response functions and neutron
fluence-to-ambient dose conversion
coefficients
Neutron Spectra/FLUKA Simulations
40°
90°

## Slide 7: Dose Measurements and Results

## Slide 8: Dose Measurements and Results
5 s spill length 2 s off
0.5 s spill length 30 s off
Total dose:
CROME PC + WENDI II = 38 µSv
FHT6020 + WENDI II = 37 µSv
CROME R + FHT 190 = 7,7 µSv
FHT6020 + FHT 192 = 8,1 µSv
Total dose:
CROME PC + WENDI II = 758 µSv
FHT6020 + WENDI II = 713 µSv
CROME R + FHT 190 = 136 µSv
FHT6020 + FHT 192 = 147 µSv

## Slide 9: Depth profiles of some residual nuclides produced in the activated Al target
Chart: Be-7 per mm per projectile
| 2.5 | 7 | 12 | 17.5 | 22.5 | 27.5 | 32.5 | 37.5 | 42.5 | 47.5 | 53 | 58 | 63 | 69 | 74 | 78.5 | 83.5 | 88.5 | 93.5 | 98.5 | 103.5 | 108.5 | 114.5 | 119.5 | 124.5 | 129.5 | 134.5 | 139.5 | 144.5 | 150.5 | 155.5 | 160.5 | 165.5 | 170.5 | 175.5 | 181.5 | 186.5 | 191.5 | 196.5 | 201.5 | 206.5 | 211.5
Be-7 2024 | 0.00055653 | 0.00090182 | 0.0011456 | 0.0014426 | 0.0015748 | 0.0018161 | 0.0019474 | 0.0019652 | 0.0019063 | 0.0017962 | 0.0017167 | 0.001739 | 0.0016994 | 0.0014947 | 0.0017621 | 0.0015623 | 0.0014923 | 0.0014625 | 0.0014875 | 0.0014149 | 0.0013204 | 0.0013053 | 0.0013214 | 0.0013336 | 0.001285 | 0.0012272 | 0.0011576 | 0.0011873 | 0.0010208 | 0.00096317 | 0.0010426 | 0.00091392 | 0.0009993 | 0.0010166 | 0.0010207 | 0.00090663 | 0.00091043 | 0.00086142 | 0.00084636 | 0.00080273 | 0.00082319 | 0.0007599
Be-7 2025 | 0.00066148 | 0.00098467 | 0.001262 | 0.001622 | 0.0018185 | 0.0019449 | 0.0019527 | 0 | 0.0017629 | 0.001682 | 0.0016358 | 0.0016351 | 0.0015826 | 0.0014584 | 0.0015147 | 0.0014622 | 0.0015841 | 0.0015244 | 0.0014495 | 0.0014496 | 0.0014437 | 0.0013778 | 0.0013483 | 0.0013244 | 0.0011239 | 0.0011312 | 0.0010767 | 0.0011746 | 0.0011639 | 0.0010116 | 0.000978 | 0.00098721 | 0.00092287 | 0.00098449 | 0.00093494 | 0.00092478 | 0.00095149 | 0.00094323 | 0 | 0 | 0 | 0
Chart: Na-22 per mm per projectile
| 2.5 | 7 | 12 | 17.5 | 22.5 | 27.5 | 32.5 | 37.5 | 42.5 | 47.5 | 53 | 58 | 63 | 69 | 74 | 78.5 | 83.5 | 88.5 | 93.5 | 98.5 | 103.5 | 108.5 | 114.5 | 119.5 | 124.5 | 129.5 | 134.5 | 139.5 | 144.5 | 150.5 | 155.5 | 160.5 | 165.5 | 170.5 | 175.5 | 181.5 | 186.5 | 191.5 | 196.5 | 201.5 | 206.5 | 211.5
Na-22 2024 | 0.00079999 | 0.0012246 | 0.0017038 | 0.0023063 | 0.002716 | 0.0030932 | 0.0033622 | 0.0033506 | 0.0037362 | 0.0036958 | 0.003725 | 0.0038615 | 0.0036146 | 0.0037035 | 0.0037843 | 0.0037322 | 0.0036444 | 0.0036517 | 0.0037758 | 0.0035824 | 0.0035069 | 0.0034567 | 0.0032926 | 0.0033326 | 0.0031933 | 0.0031546 | 0.0030007 | 0.0029424 | 0.0031187 | 0.0026862 | 0.0027335 | 0.0027581 | 0.0027225 | 0.0026796 | 0.0026518 | 0.00261 | 0.0025103 | 0.0025049 | 0.0022133 | 0.0023314 | 0.0022988 | 0.0022393
Na-22 2025 | 0.0006289 | 0.0011983 | 0.0017271 | 0.0022867 | 0.0026847 | 0.0030901 | 0.0033426 | 0.003632 | 0.0038397 | 0.0037056 | 0.0036622 | 0.0037243 | 0.0035645 | 0.0037239 | 0.003622 | 0.0034921 | 0.0035808 | 0.0035052 | 0.0034764 | 0.003383 | 0.003225 | 0.0031737 | 0.003192 | 0.0030562 | 0.0029981 | 0.0030071 | 0.0029766 | 0.002957 | 0.002704 | 0.0027577 | 0.0026761 | 0.0026663 | 0.0025963 | 0.0025236 | 0.0025243 | 0.002483 | 0.0023092 | 0.002278 | 0 | 0 | 0 | 0
Chart: Depth profile of Na-24
Five different gamma detectors were used to measure the gamma spectra of the foils
E. Mustafin

## Slide 10: Summary
Different dosimeters were tested in „classical GSI“ radiation field from 990 MeV/u uranium beam hitting the Al-target, different spill lengths and intensities were used. The irradiation was done with slow extracted beam.
The measured with different devices dose rates were compared between each other
Additional comparison of the measurements was done with the FLUKA simulation results
The Wendi-II - neutron dosimeter, FHT190/FHT192 – ion chamber, FHZ 632L-10 – proportional counter tube and new CROME Prototype showed reliable results and approved themselfs to be the perfect candidates for dose measurements at FAIR facility
The next step is to qualify the various combinations of detectors and electronics at pulsed radiation fields produced with a fast extracted beam. In this case, all possible scenarios will be covered and the measuring detector systems will be fully characterized.
Thank you very much for your attention and support during the experiment, and for providing us with the beam!