## Slide 1: Field measurements and simulations for a hybrid twin (avatar) of the Super-FRS magnets
H. Müller, S. Russenschuck

## Slide 2: Motivation
Simulations for magnet design
(„perfect magnet“)
Manufacturing of magnets
(with manufacturing errors)
Magnetic measurements
(deviations from „perfect“ magnet)
Qualification
yes/no use in Super-FRS
uncertainties/deviations remain
Refine simulations (data driven modelling)
include and weigh uncertainties in
numerical simulations
magnet production
magnetic measurements
for each magnet
Reconstruction of local field distribution
Include environmental effects like
cross-talk
beam tube
magneto-mechanical effects
powering
Model of Super-FRS string
Benefit for beam optics simulation and machine optimization
(as easier tuning of the machine)
Up to Now
Avatar (hybrid twin) approach

## Slide 3: Super-FRS: Specifications
Note: Field quality requirements depend on the position of the magnet in Super-FRS
Specification defines most stringent values
Dipole
Quadrupoles

## Slide 4: Design: Simulations from manufacturers
To check if the design meets requirements
Based on ideal material parameters (BH-curves)
Perfect magnets (no manufacturing errors)
Tolerance analysis only for some selected cases (at best)
Inevitable „model-order reduction“; not all details considered in the FEM model

## Slide 5: Magnet Qualification (dipole as example)
Polarity check
Central nominal field "search"
Central field loadline
Fringe field loadline
PCB coil calibration
Integral field loadline + homogeneity at midplane
Integral homogeneity in the GFR
PCB coil calibration
Gaussmeter/polarimeter
NMR
NMR
Hall probe
TF
TF
TF
TF
Measurement program
Verification of specifications
Translating Fluxmeter (TF)

## Slide 6: Magnet Qualification (dipole as example)
check by
Super-FRS beam optics
approval: yes/no
deviations from simulations
Standard procedure sufficient to decide on use of magnet in Super-FRS
Field quality
Load Line
Not check of specification
but necessary for magnet operation
measurements of many current levels needed

## Slide 7: Uncertainties
Simulations:
Random machine rounding errors
Systematic errors from uncertainties in (superconducting) material data
Approximation errors from FEM discretization
Inadequate knowledge of dynamic effects (hysteresis, eddy currents)
Ignorance (unrecognized systematic uncertainties), e.g., magnetostriction
Magnet construction
Random manufacturing tolerances
Systematic errors from assembly processes
Inadequate knowledge of the geometrical parameters
Ignorance (unrecognized systematic uncertainties), e.g., material parameters and environmental/operational effects (cross-talk, power converter settings)
Magnetic measurements
Random capture and read-out noise
Systematic errors from voltage dividers, calibration, integrator drift, or stray fields
Approximation errors from the projection of measurement data onto Fourier polynomials
Inadequate knowledge of the effects of environmental conditions
Ignorance (unrecognized systematic uncertainties), e.g., hysteresis in sensor response

## Slide 8: Uncertainties (dipole as example)
Field quality out of specification for
downward hysteresis branch at low currents
Due to hystersis effects not considered initially? Wrong material parameters? Integrator drift?

## Slide 9: Uncertainties (dipole as example)
Local defects in
magnetic field
Due to semi-yoke
manufacturing in
three parts

## Slide 10: Uncertainties (dipole as example)
Fringe field and cross talk
Proximity effect on fringe fields of neighboring magnets
cannot be measured with
current test facility set-up
Beam pipe
Influence of beam pipe, especially the welding seams
has not (yet) been measured

## Slide 11
Raw signal (voltages)
Processed data (fluxes)
Modelling errors
Numerical errors
Manufacturing tolerances/
errors
Capture noise
Read-out noise
Prediction
Magnetic flux density in the magnet bore
Magnetic flux
density in W
Simulated ????? Measured
Bringing together Simulations and Magnetic Measurements
Numerical model (design)
Physical object (magnet)
Calibration
Approximation
errors
Approximation
Field transducer

## Slide 12
Observation
function
Prediction
Comparable quantity
Error
metric
Quantity of interest
Expected voltages in transducer, including uncertainty
Physical state variables
Boundary sources: single and double-layer
potentials, or harmonic coefficients on trivial domains
Magnetic flux density in
trivial domain W
Inverse field problem
Generalized Field Description
Raw signal (voltages)
Processed data (fluxes)
Magnetic flux density in the magnet bore
Field transducer
Physical object (magnet)

## Slide 13
Observation
function
Prediction
Comparable quantity
Bayesian
Inference
Quantity of interest
Inverse field problem (ill-conditioned)
Field Mapping with the Moving Fluxmeter (Established: Liebsch, Fleig)
Prior
Raw signal (voltages)
Processed data (fluxes)
Magnetic flux density in the magnet bore
Field transducer
Physical object (magnet)
Magnetic flux density parameterized by stream
function on 3D domain boundary
Field Distribution

## Slide 14
Observation
function
Error
metric
Prediction
Comparable quantity
Quantity of interest
Expected voltages in transducer, including uncertainty
Physical state variables
Boundary sources: single and double-layer
potentials, or harmonic coefficients on trivial domains
Magnetic flux density in
trivial domain W
Ignorance (Unrecognized Phenomena)
Raw signal (voltages)
Processed data (fluxes)
Magnetic flux density in the magnet bore
Field transducer
Physical object (magnet)
Inadequate knowledge and ignorance

## Slide 15
Observation
function
Prediction
Comparable quantity
Quantity of interest
Expected voltages in transducer, including uncertainty
Physical state variables
Boundary sources: single and double-layer
potentials, or harmonic coefficients on trivial domains
Magnetic flux density in
trivial domain W
String of Magnets and Real-Time Machine Control
Raw signal (voltages)
Processed data (fluxes)
Magnetic flux density in the magnet bore
Field transducer
Physical object (magnet)
Power converter
control
Manufacturing
errors and features
Proximity
effects
Hysteresis
Time-trans.
effects
Field-dependent
mechanical
deformations
DX
Monitoring and
Machine Learning (GBT)
First-principle (Newton)
Data model (Kepler)

## Slide 16
Avatar
DX
Monitoring and
Machine Learning (GPT)
First-principle (Newton)
Data model (Kepler)
Design of experiment
Global/local field relation
Sensitivity analysis
Maxwellification
Interpolation of
states (powering cycles)
Empirical fits for:
Rate-dependent dynamic effects and hysteresis
Machine model
Supervised
(physics-informed)
Unsupervised
Notes: Generative pre-trained transformers

## Slide 17: Avatar
Benefits:
Identify reasons for the observed deviations of the nominal beam parameters
Help in the setting and tuning of the machine
Simulate magnet behavior at unmeasured current levels
Reduce magnetic measurements to a lower number of excitation currents
Effort:
Set-up of the Avatar model (not time or schedule sensitive)
Additional magnetic measurements required (must be done now)

## Slide 18: Box measurements with the moving fluxmeter
Required: Local field distribution
Can be calculated from field measurements on the closed surface of a box (normal components of the magnetic flux density on the boundary of a trivial domain)
Can be measured for horizontal and vertical fields
with existing equipment at CERN (longitudinal resolution <= 1 mm)
2-3 days additional measurement time needed

## Slide 19: Proposal for collaboration
Scope
GSI deploys 1 PhD student, who is working on the Avatar topic (official supervisor: Oliver Boine-Frankenheim)
CERN provides the post-processing software for the generalized field description as developed in the framework of a previous PhD.
CERN provides best-effort support for the above-mentioned PhD student
Project responsible from CERN-side: Carlo Petrone
Expenses accruing for this project will be paid from the FAIR/GSI budget code.
Necessary magnetic measurement time for the Avatar will be granted by GSI