## Slide 1
Beam Diagnostics for the FAIR Proton LINAC
RFQ Workshop at GSI
P. Forck for LOBI Department
November 20th, 2013
Outline
Diagnostics devices foreseen at pLINAC
Comments to special devices
Preliminary layout of test bench for commissioning
Interesting question for us
What are the lessons learned at LINAC4 RFQ commissioning?

## Slide 2
Beam Diagnostics for FAIR p-LINAC
Beam parameter | Device | # | Non-dest. | Remark
Current | Transformer | 9 | Yes | Dynamic transmission control
 | Faraday Cups | 3 | No | Only used as beam dump
Profile | SEM-Grid | 4 | No | Standard
Transversal emittance & halo scraping | Slit-Grid | 3 | No | Standard
Position & mean energy | BPM | 14 | Yes | Position, relative current & time-of-flight or phase
Bunch shape | Non intersecting BSM or ‘Feschenko type’ | 1 | Yes or no | Novel, under development or know technology by INR
LEBT: charge states | Wien Filter | 1 | No | From CEA/Saclay
LEBT: emittance | Allison Scanner | 1 | No | From CEA/Saclay

## Slide 3
Diagnostics Elements (except BPM)
For the p-LINAC diagnostics:
3 Faraday-Cups: basically used as beam dumps  not critical
9 ACCTs: Detector either GSI design of commercial type i.e. improved Bergoz type
DAQ: VME ADC hardware, FESA software, Java GUI,
first version for CEA/Saclay ion source commissioning
4 SEM-Grids: GSI standard, DAQ via I-to-f converter CEA/Saclay ion source comm.
Will be used for emittance measurement in connection with slits
4 Slits: For halo scraping and emittance measurement
1 Bunch Shape Monitor: see below
BPMs (not shown): see below
S
I
Faraday Cup
ACCT
Slit & SEM-Grid
Bunch Shape M.
ion source
RFQ
3 x CCH
3 x CH
to SIS18
beam dump & test bench
LEBT
Slit

## Slide 4
GSI designed ACCT:
Transformers are widely used at GSI, improved version with 2-fold bandwidth required,
possible collaboration with company Bergoz
Beam loss control via pulse-by-pulse transmission meas. control of subsequent pulse
DAQ: VME based commercial ADCs, software FESA (as GSI standard)
Current Measurements with ac-Transformers
Core radius | ri =30 mm, ro=45mm
Core thickness | 25 mm
Core material | (CoFe)70%(MoSiB)30% Vitrovac 6025
Number of windings | 2x10
Resolution | 0.2 μA for full BW
Bandwidth | 500 kHz
Droop | 0.5 % for 5 ms pulses
Max pulse length | 8 ms
Standard ACCT at GSI
H. Reeg (GSI) et al., Proc. EPAC’06

## Slide 5
Profile Measurement by SEM-Grid
SEM-Grids:
Wire spacing 0.8 to 2 mm, #wires 15 to 63 per plane, Ø0.1 mm W-Re alloy
Different sizes at LEBT and LINAC area
Multi-channel current-to-frequency conversion as GSI development (lower costs as I-U)
First version used for ion source commissioning at CEA/Saclay
Emittance: Slit-Grid method,  mounted on stepping motor feedthrough.
SEM-Grid feed-through on CF200:
I-to-f converter & digitizer ASIC

## Slide 6
Investigations for BPMs
BPMs will be the main operation tool
for position, time-of-flight & rel. current:
Button type due to simpler mechanics
Detailed CST calculation performed
0 1 2 3 4 5 6 7
time [ns]
voltage [V]
1
0
-1

## Slide 7
Investigations for BPMs
BPMs will be the main operation tool
for position, time-of-flight & rel. current:
Button type due to simpler mechanics
Detailed CST calculation performed
4 BPMs in ‘intertank’ closed to cavity
 rf pick-up from cavity
CST calculation  acceptable level
Mechanical design started
10 BPMs in regular beam pipe  mechanically less critical
 non critical
CCH2
Cavity
CCH2
Cavity
Ø20
Ø30
600 mm
100 mm
CCH1
Cavity
48 mm
28 mm
BPM button of Ø14
quadrupole

## Slide 8: Digital BPM Electronics: Under-sampling using LIBERA
Expected Advantage:
Single electronics + digital processing  amplitude = position & phase = ToF
Input: 4 plate signals & master-oscillator via under-sampling fsample = 4/11 ∙ frf = 118 MSa/s
Present status of LIBERA electronics:
General functionality shown, but some performance problems
Some problems and inconsistencies at UNILAC with frf = 108 MHz
Further test at UNILAC planned
 Applicability of concept and hardware to be proven.
Position and phase processor ‘Block diagram’
BPM signal
(left)
BPM signal
(right)
BPM signal
(top)
BPM signal
(bottom)
rf master
325.2 MHz
Filter
Filter
Filter
Filter
PLL
16 bit ADC
FPGA
Signal Processing
LAN adapter
LAN
118.3 MSa/s
clock
Filter
Filter
Filter
Filter
reference input
16 bit ADC
16 bit ADC
16 bit ADC
16 bit ADC
LIBERA from I-Tech
as Slovenian
in-kind contribution
x 4/11
Additionally:
sum signal to
20 GSa/s scope

## Slide 9
Bunch Structure using secondary Electrons: ‘Feschenko Type’
Secondary e− liberated from a wire carrying the time information.
Working principle:
insertion of a 0.1 mm wire at  10 kV
emission of secondary e− within less 1 ps
rf-deflector as ’time-to-space’ converter
detector with a thin slit
slow shift of the phase
resolution 1o < 10 ps
Status:
INR Moscow can produce such monitor
Status: Place of order nearly done
SEM: secondary electron multiplier

## Slide 10
Bunch Structure using secondary Electrons: ‘Feschenko Type’
Secondary e− liberated from a wire carrying the time information.
SEM: secondary electron multiplier
rf deflector
movable
HV wire support
electron detector
beam
Example: SNS type,
but comparable monitor for LINAC4
Alternative: Secondary electrons form residual gas
 ongoing development at GSI for UNILAC

## Slide 11
Test Bench for stepwise Commissioning
Test bench for characterization of all relevant beam parameters:
contains all monitors + magnetic spectrometer for energy measurement
+ quadrupoles and steerer
Status: Details not designed yet ! (Partly, due to ‘political’ reasons.)
Slit
BPM
BPM
ACCT
ACCT
BPM
BSM
SEM-
Grid
Cup
Slit
Slit
SEM-
Grid
Cup
Dipole

## Slide 12
Summary
Summary:
Standard devices: ACCT, SEM-Grid, Slit, Cup partly designed
Partly new technologies used e.g. SEM-Grid digitalization
New developments for BPM readout (digital ampl. & phase determination)
‘Feschenko monitor’ never used at GSI
Test bench not designed yet
What is the lesson to learn from commissioning of LINAC4 RFQ?
Could the RFQ be sufficient characterized?
Thank you for your attention.