## Slide 1
SARRP characterisation for small animal irradiation dosimetry and imaging
Dario Langellotti (University of Naples)

## Slide 2: SARRP
Where? Cave M
How it works?
CT
Small Animal Radiation Research Platform
  In Vivo and in vitro irradiation and imaging

## Slide 3: Table of content
1
2
1
Materials & methods:
SARRP
Ionization chambers & GafChromic film EBT3
2
Results:
Dose study and characterization
Instrumentation calibration and repeatability
Field uniformity for both collimated and open fields
3
Conclusions

## Slide 4: Motivation
accuracy and reproducibility of experimental irradiations.
Characterization of the Collimated Field (10 mm × 10 mm) and the Open Field
Challenges of the open field
Main Goals
Why is accurate dose characterization important?

## Slide 5: Materials & methods(1) – SARRP
Have you ever had an X-ray?
Tube voltage
Tube current
Main charateristics:
5 to 225kV
0 to 30 mA
/ 1 mm Al
but do you know how it works?
e-
Filter
Electrons emitted from the cathode
Accelerated toward a metal target
X-rays production
0.15mm Cu

## Slide 6: Material & methods (2) - SARRP
Ionization
chamber
Collimator
5.25 cm
Romeo elettrometer

## Slide 7
Ionization
Production of secondary electrons
Electric potential is applied
Ion collection
Material & methods (3) – Ionization chambers

## Slide 8
Ionization chambers
Flash Diamond
1.1 mm radius x 1 μm thickness
Semiflex
2.75 mm radius x 16.25 mm lenght
Farmer Chamber
3.05 mm radius × 23.0 mm length
Advanced Markus
2.5 mm radius × 1 mm depth
PinPoint
1 mm radius × 5 mm length
Mini PinPoint
1.45 mm radius × 5 mm length
Material & methods (4) – Ionization chambers

## Slide 9
Before irradiation
After irradiation
Polyester (125 μm )
Polyester (125 μm )
Active layer (28 μm )
Gafchromic EBT-3 [0.2 – 10] Gy
Ionization induces polymerization
Changing the optical properties of the film
High spatial resolution
lower accuracy
Material & methods (5) –GafChromic films

## Slide 10: Results (1)
Flash diamond
0.94 % deviation to Semiflex
Other chambers
Good agreement and low relative error
Experimental Repeatability
FUN FACT
Linear distribution as expected (MiniPinPoint was used)
Differences in dose measurement between 10mm x 10mm collimator and open field
  Transient effects and beam switching delays
The photoelectric effect is emphasized.
modern PinPoint
old PinPoint
6 fields were used and 3 repetitions were performed

## Slide 11: Results (2)
Deviation to deliver dose less than 2%
A copper filter provided differences in homogeneity along the two axes.
Fifth-polynomial fit
  Used for the calibration of the flm
  Correspondence between grayscale value and dose for each pixel
5 Gy
1 Gy

## Slide 12: Results (3)
17% overdose relative to the prescribed dose (5 Gy)
More than 12 Gy!!
WHY IS IT A PROBLEM?
25.4 cm
20.3 cm

## Slide 13: Conclusions (1)
Good overall agreement between delivered and prescribed dose
FlashDiamond showed the best agreement, with a deviation of only 0.94% compared with the reference dosimeter, even when 9 fields were used.
A suggestion:
Please don’t use old PinPoint chamber!!

## Slide 14: Conclusions (2)
The open-field configuration exhibited significant spatial inhomogeneities, with an approximately 17% overdose in the central region and local doses exceeding 10 Gy.
The 10 mm × 10 mm collimator provided the most suitable conditions for reproducible preclinical irradiation
Be careful when you use open field!

## Slide 15: Acknoledgment
… e grazie per l’attenzione!!!