## Slide 1: The Silicon Tracking System of the CBM experiment at FAIR
The CBM experiment, physics, technical challenges
The Silicon Tracking System
      design constraints
      layout
      tracking performance
Prototype components
Laboratory and in-beam tests
System Integration
Summary
Minni Singla, GSI
for the CBM Collaboration
HIC for FAIR Physics Day
10-04-2015

## Slide 2: The CBM experiment at FAIR
explore QCD phase diagram in the regions of high net baryon density and moderate temperatures
STS
MUCh
TRD
TOF
PSD
ECAL
RICH
MVD
Tracking
  MVD
  STS
  s.c. Dipole Magnet
PID
  RICH + TRD
  MUCH
  TOF
Event Characterization/photons:
  PSD
  ECAL

## Slide 3: Technical challenges
GEANT simulation of STS from (UrQMD) one Au+Au collision at 25 GeV/nucleon
rare probes: require 105 - 107 Au+Au reactions/sec
Avoid multiple scattering  Low material budget
Max. Fluence 1×1014neqcm-2 Radiation tolerant sensors
Rare probes:
CBM challenge

## Slide 4: Design Constraints in STS
STS
2.5°
25°
Beam
coverage
  rapidity from center-of mass to close to beam
  aperture 2.5° <  < 25°
low mass large-area detector
  high-resolution momentum determination
  track matching into MVD and RICH/MUCH
momentum resolution
  δp/p  1%
  field integral 1 Tm
  25 µm single-hit spatial resolution
material budget per station up to 1% X0
efficient hit & track reconstruction
  close to 100% hit and tracking eff.
read-out
  free-streaming fast readout
  online event selection

## Slide 5: Layout: Silicon Tracking Station
8 tracking stations
stations arranged in 4 duplets
minimizing amount of modules
double-sided micro-strip silicon sensors
a few different sensor sizes
granularity according to the hit densities
minimized material budget 0.3 -1.1% X0
read-out electronics outside the physics aperture

## Slide 6: STS performance
momentum resolution
δp/p  1%
track reconstruction efficiency
UrQMD + Geant+ CBMRoot, Au+Au @ 25 GeV/nucleon

## Slide 7: STS building blocks
Components breakdown:
  106 ladders (17 types)
  896 modules
  1220 sensors
  14144 chips
  1.8 Mio channels
Mechanical prototype module: sensor + cables + electronics
sensor
read-out cables
FEE

## Slide 8
8 tracking stations
6.2 cm × 6.2 cm
6.2 cm × 4.2 cm
6.2 cm × 2.2 cm
252 single sensors
648 sensors, daisy-chained in pairs, yielding 324 sensors (6.2 cm × 12.4 cm)
sum: 900
sum: 260
sum: 60
+ small number of sensors
3.2 cm × 2.2 cm
1.4 m
Station 1
station 8
daisy-chained pair of 6.2 cm × 6.2 cm sensors
Microstrip sensor for CBM-STS

## Slide 9: Daisy chained sensors
6 cm × 6 cm
+
2 micro cables
interconnecting the AC pads of the top and bottom sensor on the p and the n sides:
attached to the sensors via tab-bonding
Cons
complicated handling
dead region between the two sensors (due to guard rings, no overlap possible)
Step 1:
Step 2:
Alternative:
exploring with CiS
6.2 cm x 12.4 cm sensor on “one piece” of silicon
6 cm × 6 cm
6 cm × 12 cm
6 cm × 12 cm

## Slide 10: Sensor Specifications
End strips
Center strips
routing line
readout direction
9
double-sided sensors, p-n-n structure
strip lengths: 2, 4, 6 cm, 12 cm (daisy chained or not)
width: 6 cm
1024 strips per side with 58 m pitch
integrated AC-coupled read-out
stereo angle front-back sides 7.5°
n side: straight strips (00 )
routing lines on 2nd metal layer or replacement with an external micro cable on the stereo angle side
radiation tolerance up to 1 × 1014 neq/cm2
7.50

## Slide 11: low-mass read-out cables
cable
sensor
  radiation length: 0.1% X0
  two signal layers
  strip pitch 116 µm
  thickness 24 µm
  additional spacer to reduce the capacitance
  TAB bonded to sensor
  1024 channels to connect per side
  in prototypes 128 channels on each side are connected
sensor
cable

## Slide 12: Readout chip: STS-XYTER
full-size prototype
signal detection from the 2-sided sensors
fast  low noise  low power dissipation
STS-XYTER ASIC
v1.0 produced (v2.0 under design) UMC 180 nm CMOS
channels | 128, polarity +/-
noise | 900 e- at 30 pF load
ADC range | 16 fC, 5 bit
clock | 250 MHz
power | < 10 mW/channel
timestamp | < 5 ns resolution
out interface | 4(5) × 500 Mbit/s LVDS
  GBTx chip-set (CERN)
Front-end Board
Read-Out Board
Data Processing Board
time-slicing
FLES farm online event computing
8 STS-XYTER chips
1/2/5 LVDS links out separate power boards
data combining
time-stamped data

## Slide 13: Front-end board
Test assembly of FEB4
FEB4 – 4 STS-XYTERv1 ASICs (under study)
intermediate prototype with less spatial constraints
FEB8 – 8 STS-XYTERv2 ASICs
under development
challenge with high via density

## Slide 14: Quality Assurance procedures
Components
  sensors
  ASICs
  Cables
Composite objects
  modules
  ladders
wafer
prober
Keithley,
LCR meter
PC with LabView
Cleanroom @ GSI Darmstadt

## Slide 15: Irradiation studies: comparison single and double metal sensors
90Sr
End strips
Center strips
routing line
interconnection type :
    Double metallization
    external cable
sensor size (6x4, 6x6 cm2)
irradiated to 2x1014neqcm-2
SMwC: Single Metal with external cable
DM: Double metal

## Slide 16: In beam test: sensors before and after n-irradiation
COSY, 2.4 GeV protons (Dec. 2014)
Non-Irradiated, double metal
Irradiated,
double metal
n side
p side
~32 ADC
~39 ADC
~20% loss in charge collection after
neutron irradiation dose of 2x1014neqcm-2
S/N still large enough

## Slide 17: In-beam test: prototype module
COSY, 2.4 GeV protons (Dec. 2014)
2nd module prototype
with CBM05 sensor
2nd batch of the
low-mass micro-cables

## Slide 18: Tests with a laser beam
signal generation in precisely known location
test stand with focused 1060 nm infrared laser
mimic minimum ionizing particles
study e.g. charge sharing

## Slide 19: Module assembly
Inter-strip cable
chip cables
32 pcs
Sensor
FEB STS-chips
“chip-cable” to sensor bonding
“chip-cable” assembly
71 (+3) components per module
non-reworkable assembly
high-yield components essential
most work-intensive part of construction effort
assembly procedures under development

## Slide 20: System integration: stations/cabling/cooling
ROB & power supply
FEB
FEB stack
cooling plant:
bi-phase CO2 2PACL
prototype TRACI-XL: 1 kW
STS electronics total: 42 kW
prototype cooling plate

## Slide 21: The Silicon Tracking System
1632 mm
1340 mm
cooling plates
ladders

## Slide 22: Summary
STS system concept and layout has been worked out
Radiation tolerance of sensors confirmed
Prototype modules tested in beam
Module assembly and system integration in progress
Full read-out chain under preparation
Full-size STS mock-up under construction
Readiness for series production of components: end of 2016

## Slide 23: Institutions participating in CBM-STS
From Germany, Poland, Russia, Ukraine:
        Darmstadt, Germany, GSI Helmholtz Center for Heavy Ion Research GmbH
        Dubna, Russia, Joint Institute for Nuclear Research (JINR)
        Katowice, Poland, University of Silesia
        Kharkov, Ukraine, LTU Ltd * Partner Institut
        Kiev, Ukraine, Kiev Institute for Nuclear Research (KINR)
        Krakow, Poland, AGH University of Science and Technology
        Krakow, Poland, Jagiellonian University
        Moscow, Russia, Institute for Theoretical and Experimental Physics (ITEP)
        Moscow, Russia, Moscow State University
        Protvino, Russia, Institute for High Energy Physics (IHEP)
        St. Petersburg, Russia, Ioffe Physical-Technical Institute
        St. Petersburg, Russia, Khlopin Radium Institute (KRI)
        St. Petersburg, Russia, St. Petersburg State University
        Tübingen, Germany, Eberhard Karls University
        KIT, germany

## Slide 24
back up

## Slide 25: STS Technical Design Report
Approved by FAIR in Summer 2013

## Slide 26: STS Module: smallest assembled functional unit
read-out cables
10 – 50 cm
sensor
2, 4, 6, (12) cm
6 cm
front-end electronics board

## Slide 27: STS pre-series prototype sensors
Prototype | Year | Vendor | Processing | Size [cm2] | Comment
CBM05C6 | 2013 | CiS | DSDM | 6.2  6.2 | Sensor R&D GSI 2007 – 2013: CBM01 – CBM05
CBM05H4 | 2013 | Hamamatsu | DSDM + DSSM | 6.2  4.2 | 
CBM05H2 | 2013 | Hamamatsu | SSSM, 0+7.5 deg | 6.2  2.2 | BMBF-JINR funded: pre-series sensors completed
CBM06C6 | 2014 | CiS | DSSM FZ + DOFZ | 6.2 x 6.2 | 
CBM06H6 CBM06H2 | 2014 | Hamamatsu | DSSM | 6.2  6.2 6.2  2.2 | 
CBM06H6 CBM06H2 | 2015 | Hamamatsu | DSDM DSSM | 6.2  6.2 6.2  2.2 | BMBF-JINR funded: pre-series sensors in production
CBM06C6 | 2015 | CiS | DSDM | 6.2 x 6.2 | 
CBM06C4 | 2015 | CiS | DSDM | 6.2  4.2 | 
CBM06C2 | 2015 | CiS | DSDM | 6.2  2.2 | 
CBM06C12 | 2015 | CiS | DSDM | 6.2 x 12.4 | 

## Slide 28: Quality Assurance procedures
HK 2.6 Mon 16:00 M/HS1, Evgeny Lavrik
Development of optical quality assurance procedures
for the sensors
HK 21.2 Tue 15:00 M/HS1, Pavel Larionov
Quality Assurance of double-sided silicon microstrip
sensors of STS
HK 12.2 Mon 17:30 M/HS1, Anton Lymanets
Quality assurance database for the CBM Silicon
Tracking System
wafer
prober
Keithley,
LCR meter
PC with LabView
Cleanroom @ GSI Darmstadt
Notes: whats the cleanroom class??

## Slide 29: Radiation tolerance tests: 241Am
p-side
n-side
S/N for MIP > 20
59.5 keV, γ line  ~16 ke- in silicon = 2.65 fC = 116 ± 3 ADC
Notes: Calibration:
Q– [fC] = 0.54 + 0.0159 · A + 7.0 · 10–6 · A2
Q+ [fC] = 0.76 + 0.0137 · A + 9.5 · 10–6 · A2

## Slide 30: In beam set up
BEAM
STS 0: Reference station 1
STS 1: Reference station 2
STS 2: Station under test

## Slide 31: Module prototype tests
Second module prototype with CBM05 sensor and second batch of the low-mass micro-cables and n-XYTER FEB
Detailed analysis in progress
COSY @ Jülich
Notes: New prototypes were build from the new CBM05 full size sensors with low-mass micro-cables. Such modules were also under beam tests which took place in COSY Julich in December of last year. Two STS sensors were used as reference stations and middle station was the one under test having the CBM05 sensor. Modules were tested using 2.4 GeV protons and since the detailed analysis is still in progress I am able to show the beam profile as seen in reference station and ADC response from the n-side of the module.

## Slide 32: Radiation environment in STS
STATION 1
(Au+Au @ 10MHz)
  radiation hard sensors
compatible with the CBM physics program
  1 × 1013 neq/cm2 (SIS100)
  1 × 1014 neq/cm2 (SIS300)
exchange of sensors after 1 × 1014 neq/cm2

## Slide 33: Material budget in STS

## Slide 34: CBM running scenario @ SIS100
Physics case | Beam intensity 1% int. target | NIEL n/cm2/2 month | Beam time (month) | Total NIEL n/cm2
Au+Au at 2, 4, 6, 8, 10 A GeV Hadrons, e+e- | 107 | STS1 max: 5x1010 STS8 max: 1x1010 | 6 | STS1 max: 1.5x1011 STS8 max: 3x1010
C+C at 2, 4, 6, 8, 10 A GeV Hadrons, e+e- | 107 | about 20 times less than Au+Au | 2 | ---
p+A | 109 | Similar to 107 Au+Au | 1 | about 2x1010
Au+Au 10 A GeV J/ψ | 109 | STS1 max: 5x1012 STS8 max: 5x1012 | 1 | STS1 max: 2.5x1012 STS8 max: 2.5x1012
Sum SIS100 |  |  | 10 | STS1 max: 2.7x1012 STS8 max: 2.5x1012
Assumption: beam for 2 month/year x 5 years = 10 month

## Slide 35: CBM running scenario @ SIS300
Physics case | Beam intensity 1% int. target | NIEL n/cm2/2 month | Beam time (month) | Total NIEL n/cm2
Au+Au at 15, 20, 25, 30, 35 A GeV Hadrons, e+e- | 107 | STS1 max: 2x1011 STS8 max: 1x1010 | 12 | STS1 max: 1.2x1012 STS8 max: 6x1010
C+C at 15, 20, 25, 30, 35 A GeV Hadrons, e+e- | 107 | about 20 times less than Au+Au | 6 | ---
p+A | 109 | Similar to 107 Au+Au | 6 | about 6x1011
Au+Au at 15, 20, 25, 30, 35 A GeV J/ψ | 109 | STS1 max: 2x1013 STS8 max: 2x1013 | 6 | STS1 max: 6x1013 STS8 max: 6x1013
Sum SIS300 |  |  | 30 | STS1 max: 6.2x1013 STS8 max: 6.1x1013
Assumption: beam for 2 month/year x 15 years = 30 month

## Slide 36: Readout chip: STS-XYTER
full-size prototype dedicated to signal detection from the double-sided micro-strip sensors in the CBM environment
fast  low noise  low power dissipation
new w.r.t. n-XYTER architecture:
  effective two-level discriminator scheme
HK 48.56 Wed 17:00 C/Foyer, Iurii Sorokin
The STS-XYTER ASIC – a dedicated front-end
chip for the CBM Silicon Tracking System
STS-XYTER ASIC
v1.0 produced (v2.0 under design) UMC 180 nm CMOS
channels | 128, polarity +/-
noise | 900 e- at 30 pF load
ADC range | 16 fC, 5 bit
clock | 250 MHz
power | < 10 mW/channel
timestamp | < 5 ns resolution
out interface | 4(5) × 500 Mbit/s LVDS
FEB cooling box
with FEB boards
Notes: ask tomas about this slide....

## Slide 37: Cooling in STS
HK 31.5 Tue 18:00 M/HS2, Jorge Sanchez Rosado
Development of a CO2 Cooling System for the CBM
Silicon Tracking System
20 kW
20 kW
STS read-out electronics dissipates ca. 40 kW
cooling with bi-phase CO2:
  high efficiency at small spatial requirement
standard for tracker upgrades at LHC
cooperation of GSI with CERN: TRACI-XL
TRACI-XL:
1 kW test cooling system
EU-FP7 CRISP Work Package CO2 cooling
(http://www.crisp-fp7.eu/)
“2-phase accumulator controlled loop”