## Slide 1: Residual Gas Ionization Profile Monitors (IPM) @ Fermilab
4 November 2015
Beam Dynamics Meets Diagnostics
Jim Zagel

## Slide 2: Fermilab In the NOvA ERA
Proton Source and Linac
Booster Synchrotron
Muon Campus
Recycler and
Main Injector
Tevatron
n to Nova
Jim Zagel, R.Thurman-Keup | Beam Profile Measurement at Fermilab Using Residual Gas Ionization Profile Monitors

## Slide 3: Fermilab IPM Basic Types
Booster (400MeV – 8 GeV)
  Electrostatic 10KV Clearing Field)
Main Ring (8 GeV -150GeV)
  Electrostatic 30KV Clearing Field
Recycler Ultra High Vacuum (8 GeV)
  Electrostatic 30KV Clearing Field, e-11 Torr vacuum.
Main Injector Mark-II (8 GeV -150GeV)
  Permanent Magnetic Field 1KG and 10KV Clearing Field.
Main Injector Mark-III (8 GeV -150GeV)
  Adds a Control Grid and Flash Test Strip.
Tevatron (150 GeV – 1 TeV)
  Electromagnet 1 KG and 10KV Clearing Field.

## Slide 4: IPM Concept Electrostatic
Clearing Field Anode (8-30KV)
Field shaping electrodes
Wire mesh RF Shield
Microchannel Plate (MCP)
Anode strips
for Signal Pickup
Beam
(into page)
Ions
Electrons
Ionization Happens

## Slide 5: IPM Concept Magnetic Containment
Magnet with vertical B field
Cathode -10 KV
Field shaping electrodes
Electron Suppression Grid
Wire mesh gate
Microchannel Plate (MCP)
Anode strips
500 mm spacing
Beam
(into page)
Ions
Electrons
Ionization Happens

## Slide 6: Main Injector’s “Three Headed Monster”
Negative Clearing Field collecting electrons

## Slide 7
From 10/3/2003 TIPM Review
Ion vs Electron
Combined Signal from Electrons plus Secondary's
Positive Ion Collection
Electron Collection

## Slide 8
Clearing Field Polarity Reversed to Collect Ions

## Slide 9
Mark-II Secondary Suppression Screen
820Ω Resistor Should Have Been 8MΩ

## Slide 10: Residual Gas Components
RGA Scan Oct-2014 Recycler Ring 400 From 401-409 BV
13:54 | USER PEAKS | Torr
2 | H2 | 3.90E-11
17 | OH/NH3 | 5.20E-11
18 | H2O | 9.90E-11
28 | N2/CO | 9.90E-11
32 | O2 | 7.70E-11
44 | CO2 | 9.90E-11
RGA Scan June 2010 Recycler Ring 419-420
H2
C
CH4
H2O
N2/CO
O2
CO2
760 Torr = 1 Atmosphere = 101325 Pascal
10-10 Torr = 1.333x10-8 Pa
USA Labs use still Torr

## Slide 11: Residual Gas Components
RGA Scan Oct-2014 Recycler Ring 400 From 401-409 BV
13:54 | USER PEAKS | Torr
2 | H2 | 3.90E-11
17 | OH/NH3 | 5.20E-11
18 | H2O | 9.90E-11
28 | N2/CO | 9.90E-11
32 | O2 | 7.70E-11
44 | CO2 | 9.90E-11
RGA Scan June 2010 Recycler Ring 419-420
H2
C
CH4
H2O
N2/CO
O2
CO2
760 Torr = 1 Atmosphere = 101325 Pascal
10-10 Torr = 1.333x10-8 Pa
Torr was named after Evangelista Torricelli, an Italian physicist and mathematician who discovered the principle of the barometer in 1644.
(One millimeter of mercury is approximately equal to one Torr.)
From Wikipedia
USA Labs use still Torr

## Slide 12: Booster
Horizontal and Vertical co-located in Long 5
These are electrostatic instruments with a max of 8 KV

## Slide 13: Main Ring/Main Injector Electrostatic
30KV Electrostatic,
Vertical was at Q103,
Horizontal at Q102.
30 KV was shown to be sufficient to overcome space charge at 8 GeV beam injection into MI, but not at 120 GeV.

## Slide 14: Main Ring/Main Injector Electrostatic
30KV Electrostatic,
Vertical was at Q103,
Horizontal at Q102.
30 KV was shown to be sufficient to overcome space charge at 8 GeV beam injection into MI, but not at 120 GeV.
These have been functionally replaced with Mark-III IPM’s
They were removed from the Main Injector and are currently being rebuilt for use in the Booster.

## Slide 15: Main Ring/Main Injector Electrostatic
30KV Electrostatic,
Vertical was at Q103,
Horizontal at Q102.
30 KV was shown to be sufficient to overcome space charge for better than 8 GeV Beam in MI.
These have been functionally replaced with Mark-III IPM’s
and are currently being rebuilt for use in the Booster.
One concern on using 30KV electrostatic cans in Booster is beam rigidity at 400 MeV Injection energy.
Beam deflection is calculated to be:
at 400 MeV, < 88 micro-radians.
at 8 GeV Extraction, < 7 micro-radians.

## Slide 16: Main Injector Mark-II
Horizontal Measurement Permanent Magnet at Q104
Independent up and downstream +/- 25mm in horizontal plane for alignment and MCP Exposure

## Slide 17: Main Injector Mark-III Vertical
The vacuum vessel moves in the plane of the measurement, while the magnet is fixed.

## Slide 18: Recycler Mark-III Vertical

## Slide 19: Tunnel Preamp Connectors

## Slide 20: New Preamp Connector Scheme
Harting connector more reliable than Winchester

## Slide 21: Preamp Connectors

## Slide 22: Mark-II, and Tevatron Internals
Tray design for quick extraction for MCP replacement, assures accurate realignment.

## Slide 23: Mark-III Internals

## Slide 24: Anode Strip Board
Main Injector/Recycler Details
  Ceramic
  Mass terminated connectors.
    80% copper, 20% space.
  50Ω Flash Test Strip
MCP Power Pads

## Slide 25: Anode Strip Board cont’d
Beam Size determines # of Strips and Pitch
Booster
  Beam sigma 4.5 mm
  48 Strips at 1.5 mm Pitch
  1.2E12 to 4.5E12 protons
Main Injector/Recycler
  Beam sigma 4.5 - 1.5 mm
  120 Strips at 0.5 mm pitch
  MI up to 6 booster batches
  RR slip stack up to 12 batches
  RR max intensity 5E13
Tevatron
  Beam sigma 1.5 mm
  128 Strips at 0.25 mm Pitch

## Slide 26: Mark-III Control Grid

## Slide 27: Control Grid Upstairs Test R4H
Turn on
Turn off

## Slide 28: Pulsing 10KV in a Magnetic Field
Recycler Vertical showed a shoulder on the bottom, (left in screenshot.)
The Control Grid Lead had broken off the feed through.
Fixed

## Slide 29: Injection Mismatch
Main Injector:
Injection tuning study.
Poor beam line tune leads to injection oscillations in both Position and Sigma for the first 300 turns.
FFT of first 500 turns yield fractional tune

## Slide 30: Main Injector / P-Bar injection Measurement
Position oscillations for >500 turns, Sigma settles in under 300 turns.

## Slide 31: Typical Data Displays
Booster LabView Front End
Main Injector Console Application
Horizontal Vertical
Top Left trace indicates beam intensity. Top Right is raw data color intensity plot
Bottom left 2 plots – can plot sigma and position vs turn, or individual turn profiles.

## Slide 32: Main Injector Damper Turn On

## Slide 33: Tevatron
Horizontal Detector
Horizontal Correction Magnet
Measured 36 Proton and 36 Anti-Proton bunches per turn using QIE Chips in tunnel.

## Slide 34: IPM at IOTA
Re-purpose the Tevatron IPM’s to the new IOTA ring
  IOTA  Integrable Optics Test Accelerator
    Noun (from Wikipedia)
      1. the ninth letter of the Greek alphabet ( i )
      2. an extremely small amount.
  Accelerator ring designed primarily to study the behavior of a nonlinear integrable lattice
    V. Danilov and S. Nagaitsev, Phys. Rev. ST Accel. Beams 13, 084002 (2010)
  The purpose of a nonlinear integrable lattice is to make beam dynamics more robust by generating large stable tune spreads and preserving dynamic aperture
    Studies require monitoring the evolution of the beam profile on short time scales to detect beam growth and instabilities
    Ideally profile measurements every turn

## Slide 35: IPM at IOTA
2.5 MeV, 1011 protons 40 m circumference (30 MHz / 2.34 MHz modulation)
Reuse the Tevatron IPM sans magnet (G. Stancari, D. Crawford)
  Space charge is small
Vacuum needs to be 10-11 Torr
Ionization cross section much higher due to low beam energy
  M.E. Rudd et al., Phys. Rev. A 28, 3244(1983)
  Balances the very high vacuum
Expected number of ionizations per turn is ~18.
  Expected number of ionizations per bunch in MI is ~100

## Slide 36: IPM Measurement Summary
All Systems
  Turn by turn measurements.
    Turns could be averaged for any accuracy desired.
  Used for injection tuning/matching.
    Routinely used for first 500 turns to see injection oscillations.
    Position and Sigma measurements anywhere in the cycle.
  Collects up to 65K samples per channel

## Slide 37: IPM Measurement Summary 2
New Main Injector
  High Speed Digitizers 80MHz
    Multiple samples each batch for better accuracy/sample
    Allows for digital filtering of signals on A/D
  96 Channels are digitized
  Control Grid to gate off electrons for unmeasured batches
    Should significantly increase MCP life time.
  Up to 65000 samples at either
    1 Batch per revolution
    Spread across all batches for a few turns

## Slide 38: Recent References
Dirk Bartkoski, PhD Dissertation:
Analysis, Prototyping, and Design of an Ionization Profile Monitor for the Spallation Neutron Source Accumulator Ring, December 2013
https://www.researchgate.net/publication/269393015
Nuclear Instruments and Methods in Physics Research Section A Accelerators Spectrometers Detectors and Associated Equipment 11/2014; 767:379–384. DOI:10.1016/j.nima.2014.09.020
9th Ditanet Topical Workshop on Non-Invasive Beam Size Measurement
for High Brightness Proton and Heavy Ion Accelerators, April 2013
https://indico.cern.ch/event/229959/
International Beam Instrumentation Conference (IBIC):
September 2015-Melbourne, Australia: http://www.ibic2015.org
September 2014-Monterey, California: https://conf-slac.stanford.edu/ibic-2014/
Next:
September 2016-Barcelona, Spain: http://www.ibic2016.org/
Note: IBIC of interest is not the INTERNATIONAL CONFERENCE ON INDUSTRIAL BIOTECHNOLOGY

## Slide 39: Extra Material

## Slide 40: Booster Injection
One concern on using 30KV electrostatic cans in Booster;
beam rigidity @ 400MeV Injection
We start with the momentum before deflection, p0 and end with a momentum p after deflection
(1)
where q is the charge of the particle and t is the duration of the deflecting field, E. Assuming that p0 is perpendicular to E, the angle of deflection is
(2)
For small transverse deflections, or non-relativistic particles, t = l/|v0|, with |v0| being the initial velocity and l being the length of the electric field region
(3)
The initial |v0||p0| can be written in terms of kinetic energy and mass and the angle is finally
Derivation and details from R.Thurman-Keup

## Slide 41: Mark-III Vacuum Vessel

## Slide 42: Main Injector Mark-III Magnet
Magnet MIIPM001 on measurement stand for integral field map.
Similar to Mark-II but smaller.
1KG center field and half correction up and downstream,
Local 3 bump shunted, so beam sees close to zero integrated field.

## Slide 43: MI Orbit Perturbation
Measured magnet integrated field is -0.001T-m
Maximum displacement around the ring for the measured field integral is
For the Main Injector ρm ≈ 27 T-m
and the maximum β is 50, Tune, ν is 0.43,
D ≈ 0.001 m

## Slide 44: MCP Test Chamber
Facility to scan MCPs for suitability and look at areas of reduced gain

## Slide 45: MCP Test Chamber Reconfigured
Ion Pump and Ion Gauge moved away from test region
Next step is to reduce the spacing between EGN and MCP under test