## Slide 1
Measurements of small impedances
Bench measurements nowadays represent an important tool to estimate the coupling impedance of any particle accelerator device. The well-known technique based on the coaxial wire method allows to excite in the device under test a field similar to the one generated by an ultra-relativistic point charge. We discuss the basics of the coaxial wire method and review the formulae widely used to convert measured scattering parameters to longitudinal and transverse impedance data. Coaxial resonator configurations are suitable for measuring small impedances. We discuss typical measurement examples of interest for the CERN Large Hadron Collider.
Mostacci
Trieste 23-27 Settembre 2013

## Slide 2
OUTLINE
Machine impedance model
Intuition and experimental verification
Resonant coaxial wire set-up
Coaxial transmission line optimisation
(Bench) measurements of small impedances
Coaxial wire
method
Small Impedance
New machine components are designed according to coupling impedance optimization.
Check for the validity of existing machine upgrading actions.
In novel machines, high performances (e.g. high intensity, big sizes) require that coupling impedance of each device is small.
EM codes estimation needs to be guided …
for small impedance
for any impedance

## Slide 3
MOTIVATION: MACHINE IMPEDANCE MODEL
Identify main impedance contributors
Assess impedance of individual elements
Sum the impedance contributions
Longitudinal impedance model
Compute beam observables
agreement?
Identify main impedance contributors
Assess impedance of individual elements
Sum the weighted impedance contributions
Transverse impedance model
Compute beam observables
agreement?
B. Salvant, CERN, Erice 2014
Measurements of observables with beam
Machine optics
(β functions)
Measurements of observables with bench
If not (most of the time),
need to reconsider…

## Slide 4
MOTIVATION: MACHINE IMPEDANCE MODEL
Identify main impedance contributors
Assess impedance of individual elements
Sum the impedance contributions
Longitudinal impedance model
Compute beam observables
Identify main impedance contributors
Assess impedance of individual elements
Sum the weighted impedance contributions
Transverse impedance model
Compute beam observables
B. Salvant, CERN, Erice 2014
Measurements of observables with beam
Machine optics
(β functions)
Measurements of observables with bench
M. Migliorati
This talk

## Slide 5
BASIC DEFINITIONS
Coaxial wire technique: resistive beam pipes, collimators, beam pipe cross section variations, magnetic kickers ...
L. Palumbo, V. Vaccaro, M. Zobov, LNF-94/041
Impedance of most element of modern accelerators is carefully optimized (measurements and simulations).
Bench
Measurement
Resonant and non-resonant bead pull methods: accelerating and/or deflecting RF devices.
not small …
Notes: Impedance data stored in a database to derive the impedance model of the machine and eventually compute instability thresholds.

## Slide 6
COAXIAL WIRE METHOD: THE INTUITION
Ultra-relativistic
beam field
TEM mode
coax waveguide
Device Under Test
Coaxial wire
Single wire centered/displaced
Two wires
2b
2a
r

## Slide 7
BEAM vs STRETCHED WIRE: SIMULATION
Electron beam travelling on axis
Electric pulse traveling on the wire
Cavity mode excitation
Cavity mode excitation and power reflection
Courtesy of M. Migliorati

## Slide 8
THE INTUITION: TIME DOMAIN VIEW
The field “slice” of a short electric pulse on a TEM line or a short bunch leads to essentially the same image current distribution on the inner surface of the beam pipe or inner surface of the outer conductor of the coaxial line.
Due to some discontinuity, the beam is loosing momentum.
Pulse energy is lost to “fill the obstacle” (time domain view, transient view).
Beam energy is lost to “fill the obstacle”
Coupling impedance accounts for such an energy loss, explained with an E field on axis.
M. Sands, J. Rees, SLAC Report PEP-95, 1974
Due to some discontinuity, the pulse on the coaxial line is loosing amplitude.

## Slide 9
THE INTUITION: FREQUENCY DOMAIN VIEW
Due to some discontinuity, the beam is loosing momentum.
Beam energy is lost to “fill the obstacle”
Impedance accounts for such an energy loss, explained with an E field on axis.
M. Sands, J. Rees, SLAC Report PEP-95, 1974
Frequency domain view (permanent regime)
The loss of amplitude of the electric signal is compensated by the source.
Electric pulse on the wire is reflected due to the discontinuity.
The presence of the obstacle induces a reduction of the transmitted signal as well as a reflected signal. Impedance can be measured with transmission or reflection measurement.

## Slide 10
L. Vos, F. Caspers, et al., AB-Note-2003-002 MD
Goal of the experiment: to investigate the transmission properties of a coated ceramic chamber with the CERN EPA electron beam.
Coated
Chamber
Ref. Chamber
Beam Dump
Magnetic field probe
EPA beam line
Titanium coating (1.5 mm)
Thickness << skin depth
DC resistance 1 Ohm
Reference chamber
(non coated)
500 MeV electron bunch
(7e10 particles, s~ 1 ns)
EPA 1999: shielding properties
EPA 2000: effect of external shields
BEAM vs STRETCHED WIRE: EXPERIMENT
Notes: The shielding properties of thin metallic layers of finite length areimportant for LHC (fast kickers, silicon detectors).
Beam electric field can penetrate through infinitely long metallic layers of thickness much smaller than the skin depth.
Bench measurements (coaxial wire) on a kicker prototype suggested that finite thin layer have different shielding properties from infinitely long ones. Numerical (HFSS) studies predicted the shielding properties of a finite length thin layer (PAC 99).
Beam measurements assess the shielding properties but show that they can be spoiled by the addition of a second layer.
Coaxial wire bench measurements on the chambers used in the beam line confirmed later the experiment conclusions.

## Slide 11
A. Mostacci, F. Caspers, et al., CARE workshop, 2004
EPA experiment: BENCH set-up
wire diameter: 0.8 mm
Matching resistors: 240 Ohm
HF magnetic
Field probe
DUT + wire
VNA
Just like in the early days of coaxial wire techniques (Sand and Rees,1974)
The same ceramic chambers used in the beam experiment have been measured within a coaxial wire set-up, using the same magnetic field probes.
ns
50 Ohm
VNA in time domain mode (Fourier Transform)
Synthetic pulse (300 MHz BW)

## Slide 12
Bench vs beam measurement in EPA
Shallow cavity
Beam/wire axis
External shield
Field probe
Ceramic chamber
Vacuum chamber
Titanium coating
The external shield is electrically connected to the vacuum chamber.

## Slide 13
Bench vs beam measurement in EPA
Shallow cavity
Beam/wire axis
External shield
Field probe
Ceramic chamber
Vacuum chamber
Titanium coating
The external shield carries image currents and field penetrates the thin resistive (titanium) layer if this external bypass is present.

## Slide 14
LONGITUDINAL IMPEDANCE WITH VNA
Device Under Test
Vector Network Analyser
Matching
Network
Matching
Network
Transmission line properties (Time Domain Reflectometry)
Insight on the measurement set-up
Matching network computation
Single wire WITHOUT matching network
Modes identification
Wire displacement effect
Transverse impedance evaluation (from longitudinal impedance)
Single wire WITH matching network (resistors + attenuators)

## Slide 15
RELATIONS BETWEEN FORMULAS
Hahn-Pedersen formula (1978)
Improved log formula (Vaccaro, 1994)
Log formula (Walling et al, 1989)
Lumped element:
DUT length << l, pipe diameter
Sands-Rees formula (1974)
Distributed impedance
The wire method is strictly valid for frequencies below cut-off.
High frequency/long DUT
Very versatile formula, also for lumped elements with reduced accuracy.
Exact in the frame of lumped element hypothesis
Implemented in modern VNA
Notes: The improved log formula may diverge for a frequency independent lumped impedance as f->0...

## Slide 16
LHC BEAM PIPE
Pumping
holes
Weldings

## Slide 17
F. Caspers, T. Scholz, Particle Accelerators, 1995
RESONANT METHODS FOR SMALL IMPEDANCES
A reliable technique is to apply an inner conductor, which is capacitively (weakly) coupled and operates as a transmission resonator
The change in resonance frequency of this TEM line resonator (as compared to a reference measurement) indicates the imaginary part of the (wall) impedance (using the identical inner conductor), and a change in Q stands for the real part.
The capacitive coupling is immune to connector repeatability, instrument
drift or calibration uncertainties.
Weak coupling on each side is estimated via reflection measurement.
The coupling elements and other discontinuities in the set-up cause detuning too, thus increasing the uncertainty of the results of the imaginary part of the impedance.

## Slide 18
LHC BEAM PIPE: EFFECT OF THE WELDINGS
Q factor measurement for different bar configurations in a coaxial resonator.
Study of the impedance of azimuthally varying conductivity.
A. Mostacci, F. Caspers, 2002

## Slide 19
BENCH MEASUREMENT RESULTS
N steel bars
steel bar j
A. Mostacci, PRST-AB, 084402, 2005
Image current constant over the azimuth at all the relevant frequencies.
Theory
1) Tune the capacitive coupling to get symmetric coupling (monitoring S11 and S22).
2) Measure the unloaded Q for each case (need only S21 meas. for symmetric coupling).
The results are samples at the mode resonant frequencies.

## Slide 20
COAXIAL resonator methods FOR KICKERS
Frequency resolution is limited due to the length of the DUT.
High-Q, narrowband resonances may be lost.
H. Day, A Mostacci, June 2013
The losses (i.e. the Q) include both that from the beam coupling impedance as well as that due to the finite resistance of the wire. Wire contribution must be removed.
The Q of a lossy line resonator is independent from the length of the resonator.
For symmetric coupling:
T. Kroyer, F. Caspers, et al., CERN AB-Note 2007-028
Automatic meas. needed

## Slide 21
COAXIAL resonator methods FOR KICKERS
Frequency resolution is limited due to the length of the DUT.
High-Q, narrowband resonances may be lost.
T. Kroyer, F. Caspers, et al., CERN AB-Note 2007-028
H. Day, A Mostacci, June 2013
Automatic meas. needed

## Slide 22
LHC INJECTION KICKERS
H. Day , CERN 2013
Extremely low impedance can be measured !!

## Slide 23
LHC TCTP collimator
N. Biancacci et al., TCTP collimator meas., 2014

## Slide 24
LHC TCTP collimator wire Measurement
N. Biancacci et al., TCTP collimator meas., 2014
Wire (0.5mm)
LHC collimator

## Slide 25
LONG. IMPEDANCE MEASUREMENT RESULT
Wire (0.5mm)
Port 1
Port 2
No ferrite to dump transverse mode
April 2014 measurement

## Slide 26
TRANS. IMPEDANCE MEASUREMENT RESULT
Parabolic fit (with uncertainty) of Z|| vs wire displacement
Wire (0.5mm)
Port 1
Port 2
April 2014 measurement
No ferrite to dump transverse mode

## Slide 27
Coaxial transmission line
REFLECTION measurement
N. Biancacci et al., TCTP collimator meas., 2014
BPM + RF fingers
Movable jaws
Characterisation of the z-dependency of the characteristic impedance of the wire.
FFT of S parameters

## Slide 28
Measure of jaws alignment
REFLECTION measurement
N. Biancacci et al., TCTP collimator meas., 2014
FFT of S parameters

## Slide 29
N. Biancacci et al., TCTP collimator meas., 2014
Coaxial transmission line properties
Matching in impedance measurement is critical.

## Slide 30
CONCLUSION
Basis of the method
Coaxial wire method
Measurements of small impedances

## Slide 31
CONCLUSION
Basis of the method
Coaxial wire method
Resonator set-up for small impedances
Measurements of small impedances

## Slide 32
CONCLUSION
Basis of the method
Coaxial wire method
Resonator set-up for small impedances
Coaxial wire set-up optimization with TDR
Measurements of small impedances
Thank you for your attention