Results of the radiator quality test
Roland Hohler
GSI, Darmstadt
JWG University Frankfurt
PANDA Collaboration Meeting
March 2-6, 2009
at GSI

---
Heraeus bar (H3)
Delivered bars are rejections !
35mm
r = (1.2 ± 0.4) mm
front surfaces very matt
17mm
r = (1.5 ± 0.4) mm
=> reflection prob. by radius: 0.970 ± 0.007
exaggerated illustration
of the cross section
2 / 15

---
Reflection coeff. uncertainty contributions
Statistics	≈ 0.3‰ for transmission
Bulk attenuation	Λbar = 281 ± 97 m (quartz)
transmission:
Mirror correction	0.9718 (correction factor)
PD homogeneity	in this talk
Cleaning	in this talk
surface homogeneity	in this talk
Front surface loss	???
Adjustment (Brewster angle)	???
3 / 15

---
Mirror correction
brewster angle (TM)
diode
quartz bar
bar
pure beam
for
pure beam
(aluminium, 535nm)
“brewster” mirror
(Al overcoated with MgF2)
cor = 0.9718 ± 0.0003
bar
pure beam
cor. transmission:
Tcor = Tmeas cor
4 / 15

---
Homogeneity of the measurement diode
reference
(normalization
point)
(160,42)
H = 0.9997 ± 0.0003
5 / 15

---
Cleaning (Lithotec, L2)
whole bar cleaned with the “wipe & dry”-method and sidolin
defined pollution: saliva (20 - 90% loss)
all:
T = 0.9855 ± 0.0080
w/o outsider:
T = 0.9887 ± 0.0017
6 / 15

---
Transmission profile (L2, thin lateral)
17mm
OUT
IN
35mm
15 refl.
16 refl.
7 / 15

---
Bar homogeneity (L2, thin lateral)
w/o edge effects
15 reflections
T = 0.9844 ± 0.0044
16 reflections
T = 0.9836 ± 0.0044
=> main error contribution
not enough statistic to seperate between bar surface inhomogeneity and cleaning effect
8 / 15

---
Transmission profile (L2, wide lateral)
31 reflections
T = 0.9581 ± 0.0095
cleaning of the wide lateral sides is more difficult
than for the thin ones
9 / 15

---
Reflection coefficient & roughness (L2)
N =	15 refl.	16 refl.	31 refl.
T =	0.9844 ± 0.0044	0.9836 ± 0.0044	0.9581 ± 0.0095
b =	35 mm	17 mm
R =	0.99918 ± 0.00031	0.99918 ± 0.00028	0.99873 ± 0.00032
σ =	21.6 ± 4.1 Å	21.6 ± 3.8 Å	26.9 ± 3.4 Å
Λquartz = 281 ± 97 m
Θ = 55.6° (Brewster)
Lithotec specification:	20 Å
10 / 15

---
Plexiglass (acryl glass)
Dim. [W,H,L] in mm: [35,20,900]	(quartz: [35,17,800])
900mm
900mm
20mm
35mm
cooling in the
float glass process
not clear
11 / 15

---
Bulk attenuation (Plexi)
T = 0.8714 ± 0.0039
nglass = 1.5275 ± 0.0009
noil = 1.470 ± 0.017
nacryl = 1.495
=> Λacryl = 18.1 ± 1.6 m
bar
glass plates with oil
12 / 15

---
Transmission (Plexi)
No surface inhomogeneity included !
thin lateral (17 refl.):	T = 0.1216 ± 0.0003 (Θ = 56.0°)
(with glass plates, spot very smeared)
negligible loss
Θ
bar
brewster for glass plate: 56.8°
wide lateral (30 refl.):	T = 0.7473 ± 0.0003 (Θ = 56.0°)
(with glass plates)
w/o glass plates:	T = 0.6384 ± 0.0003 (Θ = 56.2°)
large difference between glass plates and w/o not clear
13 / 15

---
Reflection coefficient & roughness (Plexi)
17 refl.	30 refl.	30 refl.
with glass plates	with glass plates	w/o glass plates
Θ =	56.0°	56.0°	56.2°
R[Θ]=0.8866 ± 0.0003	0.9922 ± 0.0003	0.9870 ± 0.0003
σ =	256.4 ± 0.3 Å	67.2 ± 1.3 Å	87.3 ± 0.7 Å
only statistical error
lateral side treated with diamond needle is very rough
sides produced by float glass technique is 3-4 times rougher
than the polished Lithotec quartz bars
14 / 15

---
Summary & To-Do-list
lateral sides of the Heraeus bars are curved
main error contribution on the reflection coefficient seems
to be the surface inhomogeneity
(need more statistic to seperate it from cleaning effects)
measured roughness of a Lithotec quartz bar is consistent with
the specifications
roughness of a acryl glass bar is much larger than for
the Lithotec bar
(especially for the sides treated with a diamond needle)
further roughness determinations (Heraeus, Lithotec, Russian)
analysis of the front surface loss
15 / 15