## Slide 1: Operational Experience with the Ionization Beam Profile Monitor at COSY
22. May 2017 | Christian Böhme, I. Bekman, K. Reimers, V. Kamerdzhiev | IKP-4

## Slide 2
Supercomputer
COSY
Prof. Grünberg, Nobel Prize Physics 2007
Bilder: Forschungszentrum Jülich GmbH
Deep Brain Pacemaker

## Slide 3
COoler SYnchrotron (COSY):
184 m circumference
Protons and Deuterons
Polarized or un-polarized
1H+: 275 MeV/c - 3,65 GeV/c
Stochastic and electron cooling
2 e- cooler: 100 keV and 2 MeV
Typ. amount of stored particles: 1010
Internal experiments and 3 external beam lines
H- stripping injection
Build from 1989, commissioned 1993
Picture: Dietrich, J. and Mohos, I., PRAMANA - journal of physics 59 (2002), 1007

## Slide 4: IPM – short history
IPM using Wegde & Strip anode (decommissioned 2009)
Prototype test for FAIR
Initial Installation 2007
Readout requirements from FAIR and COSY different
  FAIR recording high speed (200 fps) and load later for processing over shared 100 Mbit network
  COSY online processing and low speed (1-10 fps), separate network (10 GBit)
Software in LabView
Optical readout using FireWire cameras (~5 m max cable length)
Triggered frame acquisition planned

## Slide 5: Initial Setup
PXI
PC
10 GBit
Tunnel
Eye
Control Room
Picture Acquisition
(PXI System from GSI)
Processing
Fire Wire
Situation Now:
FireWire Cameras are replaced with GigE Cameras
PXI System replaced by PC
Principle architecture kept
  PC in eye quite old, still running under Windows XP and Labview 8.6

## Slide 6: IPM Architecture
Vacuum
Air
Multichannel Power Supply
Ions
Electrons
-V0
+V0
UV Lamp (Calibration)
Air
Vacuum

## Slide 7: IPM Design
Picture by T. Giacomini, GSI
1: Beam
2: Camera
3: E-Field minimizing apreture

## Slide 8: Camera View

## Slide 10: Sensivity measurement using UV Lamp

## Slide 11: MCP Degredation Correction

## Slide 12
Raw data and gaussian fit
Gaussian fit parameters vs. time
Beam Current vs. time

## Slide 13: HV Control

## Slide 14: Calibration
K. Wittenburg, PKTR Note No. 70, DESY (1991)
Measure the beam width with different E-field voltages
Extrapolate to infinite voltage
In our case above ±2500 V the error in beam width is less than 5%

## Slide 15: Long Term Camera Degradation
Replacement
If the measurement software is closed the accumulated color values are saved
Black is 0, white 255
This should give an indication of dead (white) pixels
It is assumed that there is no beam picture any longer (HV off)
Single points might deviate

## Slide 16: Long Term MCP degredation

## Slide 17: Gaussian Fit
Beam parameters are mainly described by the parameters of an Gaussian fit
Gaussian fit not always correct
Upper picture: oscillating beam (exp. Time 100 ms range)
Lower P: Cooling

## Slide 18: Problems
Phosphor defects
Unsymmetrical
Effects

## Slide 19: Problems
Phosphor defects
  Reason known, solution possible
Camera defects
  Can be measured and mapped out
  Last GigE-cameras lasted ~6 years
Unsymmetrical effects
  Occur not all the time
    No reflections
  Is beam-dependent (no beam, no effect)
  Charge-up effects?

## Slide 20: Why isn‘t it fully automized yet?
HV control program
  Automatized control of the HV at channel plates, for having a signal not to low or high (damage of CP), but also not changing all the time
Vacuum control connection
  Emergency shut-of if vacuum level is to high

## Slide 21: Future Plans
Major software upgrade planned
  New control-PC should take over calculations, new software
  Publishing results over EPICS
  Possibility to integrate in CSS
  Therefore development on old system was reduced
Frame trigger pulse implementation