## Slide 1
9th March 2017
The Vacuum System of MAXIV
Beam Dynamics Meets Vacuum, Collimations and surfaces.
On behalf of the vacuum team
Eshraq Al-Dmour,
MAX IV laboratory

## Slide 2: Contents
Facility layout
Vacuum system layout.
NEG-coating R&D at CERN,
Installation procedure,
Vacuum commissioning status.
Conclusion

## Slide 3
MAX IV facility

## Slide 4
MAX IV layout
https://www.maxlab.lu.se/maxiv
3 GeV ring
(528 m)
Linac (300 m)
Short pulse facility
1.5 GeV
ring (96 m)
Conventional
design
Small aperture, all NEG-coated
Max IV - Synchrotron light source facility in Lund, Sweden.
Commissioning started:
August 2015
Commissioning started:
September 2016

## Slide 5
MAX IV, 1.5 GeV storage ring vacuum system layout

## Slide 6: 1.5 GeV storage ring layout
96m circumference
12 DBA.
12 straight sections of which, all for ID, except:
  Ach. 1: injection
  Ach. 5: RF straight Ach. 3: inj. Kicker +ID

## Slide 7
Conventional vacuum system:
  St. steel VC with lumped and distributed absorbers.
  Ion pumps, TSP and NEG strips.
1.5 GeV storage ring vacuum system
One achromat.

## Slide 8
1.5 GeV storage ring vacuum system
20 mm
Vacuum chamber during the testing.

## Slide 9
  Installation.
Assembly tables
Strong-back
1.5 GeV storage ring vacuum system

## Slide 10
  Installation.
Bakeout oven
1.5 GeV storage ring vacuum system

## Slide 11
1.5 GeV storage ring vacuum system

## Slide 12
MAX IV, 3 GeV storage ring vacuum system.

## Slide 13
3 GeV ring layout
100 MHz RF cavities
Landau cavities
Injection
Emittance measurement
Circumference 528 m,
20 achromats,
19 straight sections for IDs
Balder
2 In vacuum undulators (2 m long)
NanoMAX IVU18
BioMAX IVU18
2 EPU (4 m long, min gap 11 mm)
VERITAS EPU48
HIPPIE EPU53
1 In vac. Wiggler (2.4 m long)
BALDER IVW50
Installed Insertion Devices:
  Legend:
  Dk: Dipole kicker (S1)
  Mk: Multipole kicker (L)
  LK: Longitudinal kicker (S2)
  VP: Vertical pinger (S2)
      L=long straight,
      S=short straight,

## Slide 14
Standard vacuum chamber geometry
One achromat
Beam direction
VC5
Distributed cooling
Ribs
Cooling for corrector
Cooling for corrector
Welded bellows
Welded bellows
Chamber body
Material: OFS copper
Inside diameter: 22 mm,
Total length: 2.5 m,
Bent part
Arc length: 1 m,
Bending angle: 30,
Bending radius: 19 m.
NEG-coated.
Beam direction
Bent part
VC5 chamber

## Slide 15
General vacuum chamber geometry
Beam direction
Welded bellows
Cooling for corrector area
Ribs
Distributed cooling
Chamber body

## Slide 16
U2, VC4
Min. clearance with the iron 0.5 mm, min. clearance with the coils 2 mm.
Ø22 (ID)
Dipole
3 GeV magnet layout
One achromat
Beam direction
Photon beam
Corrector
25mm
Chamber ID 22mm
Magnet apertures Ø25mm
Sextupole
Ø25
- Total length ~26 m,
- 4.5 m straight section (L)

## Slide 17
Vacuum achromat layout
Photon beam
Electron beam
Penning gauge (S2)
Extractor gauge (S1)
In each achromat:
10 BPMs,
3 pumping ports (with ion pumps) and 1 pump in FE,
1 crotch absorber,
3 gate valves.
One achromat
Beam direction
VC1
VC2
VC3
VC4
VC5
VC6
VC7
VC8
VC9
VC10
One achromat
Length: ~26m

## Slide 18
NEG-coating R&D at CERN
Define and perform initial surface treatment of OFS copper substrate.
Validate compatibility of NEG-coating (adhesion, thickness, activation behavior):
a). on etched OFS copper.
b). on wire-eroded surfaces and used brazing alloys.
To validate the coating feasibility 3 main stages of NEG (Ti, Zr, V) coating validation by magnetron sputtering in collaboration with CERN were undertaken. (R&D duration ~2 years).
3. NEG-coating validation of compact vacuum chamber geometries:
a). Coating and testing of small diameter, bent tubes.
b). Establish coating procedure/technology and coat chambers
of complex geometry.

## Slide 19
Ø25
5
7
Ø22
3
Chamber exit
Chamber entrance
Chamber exit
Chamber entrance
0.75 m
Aperture limiting sextupole
NEG coating: R&D at CERN
Example: Develop coating procedure for chambers with small antechamber –(vertical aperture from 5 to 7 mm).
Prototype made at CERN in two halves to be able to inspect the coating quality.

## Slide 20: NEG-coating series production
1. Standard bent vacuum chambers (VC4) - 1.50 and 30 bends,
2. Straight vacuum (VC10) chambers,
3. Special/complicated vacuum chambers.
30
2.8 m (VC4)
2 m (VC10B)
Electron beam
Photon beam
VC2
BPM
17.2 mm
Ø22
Ø22 mm
One achromat
Beam direction
3 GeV - 1 achromat (~26m)
VC4
VC2
Industry (70% length wise)
Collaboration with ESRF (15%)
Collaboration with CERN (15%)
Main vacuum chamber types:
VC10
Photon beam

## Slide 21: Installation procedure
Ring installation was tested and rehearsed by installing and activating 1 mockup achromat in summer 2014.
Actual installation started in November 2014, ended June 2015
7 magnet blocks on concrete girders
Strongback
7 magnet blocks on concrete girders
Installation done with help from Budker Institute

## Slide 22: Installation procedure
Assembly in-situ (above magnets),
Pump down and testing,
Lifting,
Bake out (1 day), activation (1 day),

## Slide 23: Installation procedure
Lowering to the bottom magnet half,
Installation of final equipment (supports, BPM cables),
Lowering to magnet block.

## Slide 24: Coating non-conformities
All the chambers were inspected at site before installation.
Peeling-off at RF fingers and
Copper end piece
Observed peeling-off:
At RF fingers Cu-Be insert and Cu end piece. RF fingers and Cu end were not shielded properly during coating. Solution: new pieces ordered and replaced (without coating).
Uncoated areas:
Few cm^2 uncoated, in complex chambers.
Uncoated areas
Peeling-off
Peeling-off at the edge of stainless VC. Chamber not aproved for installation.
peeling-off

## Slide 25
Commissioning
3 GeV ring

## Slide 26
Commissioning progress
3 GeV storage ring commissioning started in August 2015
Accumulated beam dose:
140 Ah.
1st shutdown March 2016:
2 in-vacuum undulators,
2nd shutdown August 2016:
2 EPU chambers (8x36mm),
In-vacum wiggler.
Average base pressure:
Gauges 2e-10 mbar,
Ion pumps in 8e-11 mbar range.
First shut down
second shut down

## Slide 27: Pressure vs beam current vs dose
Beam direction
Chart: I [mA]
P S1 EXTR. Gauges (16 Ah) | 3.8e-10 | 4.24e-10 | 5.08e-10 | 5.79e-10 | 6.48e-10 | 7.26e-10 | 8.07e-10 | 8.93e-10 | 9.72e-10 | 1.05e-09 | 1.16e-09 | 1.25e-09 | 1.34e-09 | 1.411e-09
Pressures at dose 16 Ah and 95 Ah, during beam ramp up. (pressure recorded by extractor gauge at not NEG coated crotch absorber in S1)
Extractor gauge

## Slide 28: Pressure in one achromat vs beam current
Beam direction
Pressures at dose 95 Ah, during beam ramp up, recorded at different positions along achromat (L, S1, S2)
Chart: Beam Current [mA]
P L IP | 8.6e-11 | 9.4125e-11 | 1.2081e-10 | 1.4513e-10 | 1.7125e-10 | 1.9937e-10 | 2.2375e-10 | 2.4687e-10 | 2.725e-10 | 2.9875e-10 | 3.225e-10 | 3.45e-10 | 3.7188e-10 | 3.9625e-10 | 4.2125e-10 | 4.4625e-10 | 4.7063e-10 | 4.9312e-10 | 5.1e-10 | 5.0687e-10 | 5.2188e-10 | 5.4313e-10 | 5.6625e-10 | 5.8875e-10 | 6.0875e-10 | 6.3187e-10 | 6.7125e-10 | 6.55e-10 | 6.4812e-10
Legend:
Pressure at S1 (ion pump)*
Pressure at S1 (extractor gauge)
Pressure at S2 (penning gauge)
Pressure at S2 (ion pump)
Pressure at L (ion pump)
* Pressure at S1 (Ion pump) probably due to electrical effect

## Slide 29: Vacuum conditioning curve
Beam direction
Normalized average pressure vs beam dose,
recorded in S1 uncoated crotch absorber, and S2 NEG-coated
Legend:
Pressure at S1 (extractor gauge)
Pressure at S2 (penning gauge)
Chart: Accumulated Dose [A∙h]
 | 1.0778e-10 | 5.4758e-12
 | 4.9916e-10 | 4.6614e-12
Extractor gauge
Penning gauge

## Slide 30: Vacuum conditioning curve
ALBA: Raquel Monge, privet communication.
Diamond: M P Cox et al, Commissioning of the diamond light source storage ring vacuum system, Journal of Physics: Conference Series 100 (2008) 092011
Soleil: J.C. Besson, et al COMMISSIONING & OPERATION OF SOLEIL, WAO 2007. PSI - Scientific and Technical Report 2003 / Volume VI
SLS: L. Schulz et al, STATUS REPORT OF THE SLS STORAGE RING VACUUM SYSTEM: EXPERIENCE AFTER TWO YEARS OF OPERATION
ASP: E. Al-Dmour, VACUUM PERFORMANCE IN THE MOST RECENT THIRD GENERATION SYNCHROTRON LIGHT SOURCES, EPAC08.
Chart: normalized average pressure vs acc. dose
| 1 | 10 | 100
ALBA | 2e-10 | 6e-11 | 1e-11
Diamond | 1.5e-10 | 4e-11 | 7e-12
Soleil | 2e-10 | 5e-11 | 1e-11
ASP | 1e-10 | 5e-11 | 1.5e-11
MAXIV | 1.5e-10 | 3e-11 | 6e-12
SLS | 2e-10 | 6e-11 | 1e-11

## Slide 31: Residual gas spectrum
Residual gas spectrum at 140 mA beam current
* Mass 16 and 19 is most likely due to intrinsic degassing of the gas spectrometer.
Main residual gasses:
Hydrogen 85.33 %,
Carbon monoxide 9.39 %
Carbon dioxide 0.42 %
Methane 2.16-1.17 %

## Slide 32: Lifetime evolution
Maximum stored beam current 198 mA

## Slide 33: Scraper measurements
Mean lifetime vs vertical scraper distance from the beam center (Done at beam dose 140 Ah)
Total average pressure along the beam path
(contribution from all gasses based on the RGA spectra Z2 =5.77)
At 70 mA (dose 140 Ah):
Elastic beam lifetime τelastic = 163 h
Inelastic beam lifetime τinelastic = 104 h
Touschek lifetime τTouschek = 63 h
Pressure as seen by the beam 7.3e-09 mbar
Scraper position from beam center, y (mm)
Mean lifetime, τ (h)

## Slide 34
Commissioning
1.5 GeV ring

## Slide 35
Commissioning progress
1.5 GeV storage ring commissioning started in Sep. 2016

## Slide 36
Chart: ion pumps inside magnet blocks. (mbar/mA)
| 0.59 | 0.11700000000000001 | 0.19 | 0.25 | 0.39 | 0.47 | 0.55000000000000004 | 0.61 | 0.63 | 1.79 | 3.07 | 4.3600000000000003 | 5.07 | 6.17 | 6.9 | 7.3 | 7.7 | 8.4 | 9.93 | 10.210000000000001 | 10.86 | 11.6 | 13.48 | 13.05 | 17.989999999999998 | 27.18 | 30.57 | 31.4 | 37.049999999999997 | 48.27
average IPN01 and 02 | 7.263e-10 | 6.7736e-10 | 5.024e-10 | 9.3167e-10 | 2.1316e-10 | 3.2158e-10 | 2.2385e-10 | 1.4884e-10 | 7.2215e-11 | 1.0182e-10 | 7.513e-11 | 1.248e-10 | 1.4576e-10 | 7.3125e-11 | 2.2712e-10 | 2.9479e-10 | 5.0116e-11 | 2.9959e-11 | 2.2721e-11 | 1.4612e-11 | 2.439e-11 | 1.3455e-11 | 1.9532e-11 | 1.2261e-11 | 7.7892e-12
Chart: ion pumps in the long straight (mbar/mA)
| 0.59 | 0.11700000000000001 | 0.19 | 0.25 | 0.39 | 0.47 | 0.55000000000000004 | 0.61 | 0.63 | 1.79 | 3.07 | 4.3600000000000003 | 5.07 | 6.17 | 6.9 | 7.3 | 7.7 | 8.4 | 9.93 | 10.210000000000001 | 10.86 | 11.6 | 13.48 | 13.05 | 17.989999999999998 | 27.18 | 30.57 | 31.4 | 37.049999999999997
LS pressure | 5.5263e-10 | 4.1032e-10 | 4.0682e-10 | 3.8077e-10 | 3.1079e-10 | 2.8431e-10 | 2.6427e-10 | 3.4146e-10 | 2.6091e-10 | 4.1391e-10 | 2.7598e-10
Vacuum conditioning curve

## Slide 37
Lifetime evolution

## Slide 38
Beam direction
Problems
RF cavity (S2) venting due to broken high power feedthrough during conditioining (with closed valves). Now cavity is removed from the ring and dummy chamber placed as could not be run with high power anymore. Now awaits conditioning outside the ring. Also pick up loops ceramic parts had leaks.
e-
Photon beam
Hot spot
Hot spots in proximity of crotch absorber (S1), and in the slit between chamber and antechamber, miss-positioning of the crotch chamber, chamber deformation.

## Slide 39
Conclusion
MAXIV 3GeV ring vacuum design, production and installation was a challange, it is the first fully NEG coated machine.
NEG coating increased the complexity of the design, production and the installation.
R&D in NEG coating was essential in making a successful production and performance of the chambers.
MAXIV 3GeV ring is under commissioning, with the first users are expecting on April 2017.
The vacuum conditioning is following similar behavior to that of similar machines which used conventional vacuum design.

## Slide 40
Thanks for your attention
Acknowledgments:
Jonny Ahlbäck, Marek Grabski, Chiara Pasquino (CERN), Esa Paju, Michael Gilg, Jens Sundberg, Åke Andersson, Pedro F. Tavares.
Collaborators: CERN, ALBA, ESRF, BINP, Soleil.