## Slide 1: Ion Pump design for improved speed at low pressure
M.Audi , C.Paolini - Agilent Technologies
Beam Dynamics meets Vacuum
KIT , March 2017

## Slide 2: Scope of the work
Ion pumps are normally used at low pressure , 10-8 mbar ( 10-6 Pa ) and below.
Existing ion pumps have maximum pumping speed in the 10-6 mbar (10-4 Pa ) range.
Analyze influence of design parameters on pumping speed of an ion pump.
Experimental test of elements with different parameteres in different operating conditions.
Goal is to optimize these parameters in order to :
-Improve the overall pumping speed
-More specifically, improve it at lower pressures.

## Slide 3
Ion Pumps main pumping principle ( for getterable gases )
Penning Cell , crossed B and V act as an electron trap
Electrons ionizes gas molecules through collisions
Ions are created in order to bombard the Titanium cathode
Even if some ions diffuse into the cathode, this is not the main pumping effect
Ion Bombardment causes Sputtering of chemically active Titanium film on the anode
Neutral Gas molecules stick to the Titanium film (chemisorption) and are buried in the anode
This is the main pumping effect and the only stable one
Ion pumps do not pump ions by diffusion into the cathode , but pump neutral molecules by chemisorption by the sputtered cathode material on the anode

## Slide 4: Ion Pump pumping speed
The amount of gas pumped is proportional to the amount of sputtered cathode material.
This is proportional to the number of ions bombarding the cathode with sufficient energy to create sputtering .
Q ( mbar l / s ) = h I ( A )
S (l/s) = Q (mbar l / s) / P ( mbar ) = h I / P
I/P , discharge intensity , must be maximized to maximize S
I/P = f ( V , B , cell d , .........)
I = k P n where usually 1< n < 1.2
I/P , and therefore S , increase with p ( up to 10 – 5 mbar )
n close to 1 to obtain constant I/P ( I/P = k ) and therefore S
n depends again on V , B , cell d

## Slide 5: Discharge intensity in Low Magnetic Field and High Magnetic Field modes
Low pressure
High pressure
P (mbar) | @ 7 kV | @ 5 kV | @ 3 kV
1,00E-06 | 1256 | 1061 | 823
1,00E-07 | 1409 | 1191 | 922
1,00E-08 | 1580 | 1336 | 1035
1,00E-09 | 1774 | 1499 | 1161
20 mm dia cell
Notes: The highest ion pumping speed is achieved in HMF mode and also in LMF mode the speed increases with the magnetic field strength (proportional to B^2).
This is a typical diagram of the discharge intensity dependence on the magnetic field at constant pressure.
The branch AO is characteristic for the LMF mode. At low pressures (p <1E-7 Torr) the HMF mode (branch OC) shows an almost constant discharge intensity. For high pressure (p>1°-7 Torr) the HMF mode is given by the curve OD.
After this introduction, we easily understand how complex is the physics of an ion pump and how difficult can be to predict (and sometimes also to measure) the pumping speed of an ion pump.

## Slide 6: I/P measurements - Experimental set up
Roughing Valve
Vacuum chamber: UHV gauge, variable leak valve, independent ion pump
“special” ion pump body with HV feedthrough
Electromagnet poles
Ion Pump element to be tested

## Slide 7: Fixed Cell Diameter , variable B
Chart: Pressure
| 2.4E-9 | 3.4999999999999999E-9 | 6.6999999999999996E-9 | 9.6999999999999992E-9 | 2.6000000000000001E-8 | 6.1999999999999999E-8 | 9.5000000000000004E-8 | 2.6E-7 | 5.7000000000000005E-7 | 8.6000000000000002E-7 | 2.3E-6 | 5.8000000000000004E-6 | 9.0999999999999993E-6 | 1.5E-5 | 2.0999999999999999E-5
7 kV - 1000 Gauss - 16 mm | 75 | 88.5714 | 101.4925 | 113.4021 | 165.3846 | 209.6774 | 221.0526 | 280.7692 | 298.2456 | 325.5814 | 386.9565 | 413.7931 | 450.5495 | 466.6667 | 466.6667
7 kV - 1200 Gauss - 16 mm | 125 | 157.1429 | 156.25 | 173.4694 | 234.375 | 247.4227 | 344.8276 | 388.8889 | 402.2989 | 444.4444 | 500 | 526.3158 | 568.75
7 kV - 1400 Gauss - 16 mm | 140 | 226.6667 | 278.6885 | 304.3478 | 366.6667 | 383.3333 | 466.6667 | 491.8033 | 505.4945 | 645.1613 | 622.9508 | 588.8889 | 573.3333
7 kV - 1600 Gauss - 16 mm | 170 | 267.8571 | 317.4603 | 347.3684 | 406.25 | 380.9524 | 440 | 433.3333 | 475.4098 | 489.1304 | 535.7143 | 557.377 | 554.3478 | 573.3333
7 kV - 2000 Gauss - 16 mm | 223.0769 | 323.5294 | 384.6154 | 410 | 424.2424 | 387.0968 | 482.3529 | 466.6667 | 491.8033 | 510.6383 | 541.6667 | 533.3333 | 539.3258 | 550
7 kV - 2200 Gauss - 16 mm | 188.2353 | 343.75 | 409.8361 | 430.1075 | 468.75 | 423.7288 | 525 | 500 | 523.8095 | 510.8696 | 517.2414 | 534.4828 | 521.2766 | 526.6667 | 500

## Slide 8: Fixed Cell Diameter, variable B
Chart: Pressure (mbar)
| 3.3999999999999998E-9 | 5.8999999999999999E-9 | 8.7999999999999994E-9 | 2.4999999999999999E-8 | 6.1999999999999999E-8 | 8.9000000000000003E-8 | 2.9999999999999999E-7 | 6.0999999999999998E-7 | 9.2999999999999999E-7 | 2.7999999999999999E-6 | 5.8000000000000004E-6 | 8.8999999999999995E-6 | 1.5E-5
7 kV - 1000 Gauss - 20 mm | 158.8235 | 169.4915 | 193.1818 | 248 | 258.0645 | 258.427 | 333.3333 | 377.0492 | 397.8495 | 428.5714 | 465.5172 | 483.1461 | 540
7 kV - 1200 Gauss - 20 mm | 140.7407 | 146.875 | 193.5484 | 211.1111 | 273.3333 | 288.1356 | 274.7253 | 344.8276 | 362.069 | 370.7865 | 392.8571 | 423.7288 | 465.9091 | 473.3333
7 kV - 1400 Gauss - 20 mm | 266.6667 | 290.3226 | 318.6813 | 375 | 383.3333 | 466.6667 | 465.5172 | 488.8889 | 533.3333 | 550 | 516.8539 | 540
7 kV - 1600 Gauss - 20 mm | 365.7 | 380.1 | 377.7778 | 448.2759 | 508.3333 | 533.3333 | 542.3729 | 538.4615 | 533.3333 | 550 | 560.4396 | 566.6667
7 kV - 2000 Gauss - 20 mm | 479.9 | 520 | 516.6667 | 500 | 533.3333 | 517.2414 | 532.2581 | 505.618 | 500 | 491.8033 | 460.6742 | 453.3333 | 450
7 kV - 2200 Gauss - 20 mm | 444.4444 | 484.8485 | 516.6667 | 510.8696 | 517.2414 | 500 | 500 | 465.5172 | 460.6742 | 433.3333 | 428.5714 | 433.3333

## Slide 9: Fixed Cell Diameter, variable B
Chart: Pressure (mbar)
| 1.8E-9 | 3E-9 | 6.1E-9 | 9.1999999999999997E-9 | 2.9999999999999997E-8 | 5.9999999999999995E-8 | 9.2000000000000003E-8 | 3.1E-7 | 5.9999999999999997E-7 | 8.8000000000000004E-7 | 2.7999999999999999E-6 | 6.1E-6 | 8.8999999999999995E-6 | 1.4E-5
7 kV - 1000 Gauss - 24 mm | 200 | 223.3333 | 262.2951 | 282.6087 | 400 | 450 | 467.3913 | 516.129 | 550 | 556.8182 | 571.4286 | 573.7705 | 561.7978 | 585.7143
7 kV - 1200 Gauss - 24 mm | 535.7143 | 486.4865 | 491.8033 | 494.5055 | 586.2069 | 583.3333 | 584.2697 | 566.6667 | 590.1639 | 573.0337 | 571.4286 | 602.9412 | 597.7011 | 620
7 kV - 1400 Gauss - 24 mm | 617.6471 | 633.3333 | 695.6522 | 633.3333 | 606.5574 | 586.9565 | 566.6667 | 573.7705 | 544.4444 | 500 | 468.75 | 494.382 | 521.4286
7 kV - 1600 Gauss - 24 mm | 629.6296 | 636.3636 | 633.3333 | 612.9032 | 593.2203 | 585.1064 | 562.5 | 532.2581 | 516.8539 | 451.6129 | 460.6742 | 473.3333
7 kV - 1800 Gauss - 24 mm | 628.5714 | 633.3333 | 600 | 554.3478 | 566.6667 | 559.1398 | 517.2414 | 516.129 | 494.5055 | 482.7586 | 466.6667 | 494.5055 | 506.6667
7 kV - 2000 Gauss - 24 mm | 636.3636 | 633.3333 | 616.6667 | 560.4396 | 566.6667 | 558.1395 | 571.4286 | 500 | 466.6667 | 473.1183 | 451.6129 | 467.7419 | 545.4545 | 525 | 528.5714
7 kV - 2200 Gauss - 24 mm | 630.7692 | 566.6667 | 550 | 544.4444 | 562.5 | 540.9836 | 527.4725 | 466.6667 | 442.623 | 411.1111 | 437.5 | 461.5385 | 488.8889 | 492.8571

## Slide 10: Fixed B, variable Cell diameter
Chart: Pressure (mbar)
| 5.4999999999999996E-9 | 1.4999999999999999E-8 | 3.2999999999999998E-8 | 6.4000000000000004E-8 | 9.2000000000000003E-8 | 3.2000000000000001E-7 | 6.0999999999999998E-7 | 9.2999999999999999E-7 | 2.7999999999999999E-6 | 6.0000000000000002E-6 | 9.3999999999999998E-6 | 1.2999999999999999E-5 | 1.7E-5
7 kV - 1000 Gauss - 12 mm | 25.4545 | 22.6667 | 25.1515 | 26.5625 | 27.1739 | 31.25 | 39.3443 | 43.0108 | 64.2857 | 75 | 96.8085 | 138.4615 | 182.3529
7 kV - 1000 Gauss - 14 mm | 22.2222 | 28 | 40 | 64 | 69.1489 | 112.5 | 136.6667 | 139.7849 | 166.6667 | 183.3333 | 200 | 254.8387 | 271.1864 | 296.7033 | 322.2222
7 kV - 1000 Gauss - 16 mm | 75 | 88.5714 | 101.4925 | 113.4021 | 165.3846 | 209.6774 | 221.0526 | 280.7692 | 298.2456 | 325.5814 | 386.9565 | 413.7931 | 450.5495 | 466.6667 | 466.6667
7 kV - 1000 Gauss - 20 mm | 105.2632 | 158.8235 | 169.4915 | 193.1818 | 248 | 258.0645 | 258.427 | 333.3333 | 377.0492 | 397.8495 | 428.5714 | 465.5172 | 483.1461 | 540
7 kV - 1000 Gauss - 24 mm | 200 | 223.3333 | 262.2951 | 282.6087 | 400 | 450 | 467.3913 | 516.129 | 550 | 556.8182 | 571.4286 | 573.7705 | 561.7978 | 585.7143
7 kV - 1000 Gauss - 28 mm | 294.1176 | 319.1489 | 320.7547 | 358.209 | 460 | 488.3721 | 520 | 540 | 515.1515 | 522.3881 | 520 | 526.3158 | 538.4615 | 533.3333 | 528.5714

## Slide 11: Fixed B, variable Cell diameter
Chart: Pressure (mbar)
| 4.2000000000000004E-9 | 6.6000000000000004E-9 | 1E-8 | 2.0999999999999999E-8 | 2.9000000000000002E-8 | 5.7000000000000001E-8 | 9.2999999999999999E-8 | 2.9999999999999999E-7 | 5.8999999999999996E-7 | 9.9999999999999995E-7 | 3.7000000000000002E-6 | 6.6000000000000003E-6 | 9.7999999999999993E-6 | 1.5999999999999999E-5
7 kV - 1200 Gauss - 12 mm | 28.5714 | 24.2424 | 24 | 30 | 32.069 | 38.5965 | 41.9355 | 63.3333 | 72.8814 | 100 | 200 | 212.1212 | 224.4898 | 237.5
7 kV - 1200 Gauss - 14 mm | 60 | 80.5556 | 84.4828 | 102.1978 | 153.125 | 180.3279 | 200 | 233.3333 | 254.2373 | 269.6629 | 320.6897 | 338.7097 | 355.5556 | 394.4444
7 kV - 1200 Gauss - 16 mm | 125 | 157.1429 | 156.25 | 173.4694 | 234.375 | 247.4227 | 344.8276 | 388.8889 | 402.2989 | 444.4444 | 500 | 526.3158 | 568.75
7 kV - 1200 Gauss - 20 mm | 140.7407 | 146.875 | 193.5484 | 211.1111 | 273.3333 | 288.1356 | 274.7253 | 344.8276 | 362.069 | 370.7865 | 392.8571 | 423.7288 | 465.9091 | 473.3333
7 kV - 1200 Gauss - 24 mm | 535.7143 | 486.4865 | 491.8033 | 494.5055 | 586.2069 | 583.3333 | 584.2697 | 566.6667 | 590.1639 | 573.0337 | 571.4286 | 602.9412 | 597.7011 | 620
7 kV - 1200 Gauss - 28 mm | 547.0588 | 538.4615 | 548.3871 | 540.5405 | 522.3881 | 500 | 454.5455 | 452.8302 | 449.4382 | 466.6667 | 437.5 | 435.4839 | 462.3656 | 471.4286

## Slide 12: Fixed B, variable Cell diameter
Chart: Pressure (mbar)
| 6.2000000000000001E-9 | 8.9000000000000003E-9 | 3.5999999999999998E-8 | 6.1000000000000004E-8 | 9.9999999999999995E-8 | 2.7000000000000001E-7 | 6.5000000000000002E-7 | 9.0999999999999997E-7 | 3.0000000000000001E-6 | 6.0000000000000002E-6 | 8.6999999999999997E-6 | 2.0000000000000002E-5
7 kV - 1600 Gauss - 12 mm | 82.2581 | 108.9888 | 169.4444 | 190 | 214.8148 | 230.7692 | 252.7473 | 303.3333 | 316.6667 | 333.3333 | 330
7 kV - 1600 Gauss - 14 mm | 140.9091 | 193.9394 | 222.2222 | 303.0303 | 317.4603 | 329.6703 | 344.8276 | 365.0794 | 369.5652 | 400 | 451.6129 | 494.5055 | 537.5
7 kV - 1600 Gauss - 16 mm | 170 | 267.8571 | 317.4603 | 347.3684 | 406.25 | 380.9524 | 440 | 433.3333 | 475.4098 | 489.1304 | 535.7143 | 557.377 | 554.3478 | 573.3333
7 kV - 1600 Gauss - 20 mm | 365.7 | 380.1 | 377.7778 | 448.2759 | 508.3333 | 533.3333 | 542.3729 | 538.4615 | 533.3333 | 550 | 560.4396 | 566.6667
7 kV - 1600 Gauss - 24 mm | 629.6296 | 636.3636 | 633.3333 | 612.9032 | 593.2203 | 585.1064 | 562.5 | 532.2581 | 516.8539 | 451.6129 | 460.6742 | 473.3333
7 kV - 1600 Gauss - 28 mm | 471.4286 | 486.4865 | 482.1429 | 473.1183 | 500 | 492.7536 | 490 | 454.5455 | 450 | 454.5455 | 433.3333 | 444.4444 | 457.1429

## Slide 13: Pump Parameter Design
Dia , mm | Area , mm2 | No of Cells
16 | 201 | 75
20 | 314 | 48
24 | 452 | 33
28 | 615 | 24
Optimum Magnetic Field for I/P 1400 / 1600 Gauss
Present Ion Pumps work with 1100/1200 Gauss, obtainable with commercially available magnets and conventional design
Optimum cell diameter 20 to 24 mm
We have to work on Magnetic Field
Present Ion Pumps have cells from 15 to 25 mm diameter

## Slide 14: Ion Pump Magnetic field
Conventional, 150 L/S
New design

## Slide 15: Ion Pump Magnetic field
Conventional
New design
Red 900 G
Orange 1000 G
Yellow 1250 G
Grey > 1500 G

## Slide 16: Pumping speed after saturation , DIODE , N2
Chart: P (mbar)
| 9.9999999999999994E-12 | 1.5848931899999999E-11 | 2.5118864237083764E-11 | 3.9810716869888603E-11 | 6.3095734056104572E-11 | 9.9999999223571226E-11 | 1.5848931776944333E-10 | 2.5118864042053669E-10 | 3.9810716560786745E-10 | 6.3095733566211132E-10 | 9.9999998447142445E-10 | 1.5848931653888663E-9 | 2.5118863847023585E-9 | 3.9810716251684887E-9 | 6.3095733076317706E-9 | 9.9999997670713677E-9 | 1.5848931530832997E-8 | 2.5118863651993497E-8 | 3.9810715942583025E-8 | 6.3095732586424274E-8 | 9.9999996894284913E-8 | 1.5848931407777333E-7 | 2.5118863456963409E-7 | 3.9810715633481165E-7 | 6.3095732096530845E-7 | 9.9999996117856166E-7 | 1.5848931284721666E-6 | 2.5118863261933325E-6 | 3.9810715324379325E-6 | 6.3095731606637424E-6 | 9.9999995341427422E-6 | 1.5999999999999999E-5 | 2.5118863066903243E-5 | 3.9810715015277468E-5 | 6.3095731116744001E-5 | 9.9999994564998679E-5
150 D, N2 | 55 | 62 | 72.605 | 77.3283 | 83.5392 | 83 | 88 | 94 | 101.5989 | 104.5281 | 108.8802 | 111.754 | 114.6062 | 116.0242 | 117.4368 | 120.9453 | 123.0347 | 124.4211 | 125.1124 | 124.4211 | 121.6431 | 127.1785 | 133.9831 | 140.6636 | 147.2233 | 149.8141 | 148.5211 | 141.9851 | 128.5495 | 107.435 | 77.3283 | 61.3611 | 49.7942 | 41.3264 | 32.6813 | 27.4067
200 D - 3 kV, N2 | 82 | 91 | 112 | 139.167 | 160.8414 | 159.6615 | 152.9327
200 D - 5 kV, N2 | 148.8353 | 167.8349 | 178.9031 | 180.5669 | 173.7286
200 D - 7 kV, N2 | 145.7305 | 174.7159 | 198.8167 | 194.5529 | 176.6612 | 173.1478 | 163.9821
NEW VIP200
Convent VIP150

## Slide 17: Vacuum Firing processed
Vacuum Furnace
@ 950°C , 10 E-5 mbar
VIP 200 bodies
Pump internal components
Notes: Decided to applied the vacuum firing to all surface in the ion pump that are exposed to vacuum in order to reduce order of magnitudes the outgassing rate.
Vacuum furnaces is in the 10-5 mbar range @ 950°C
You can see the pump bodies and some internal parts of the pumps (in this case some anodes of the elements) ready to be «treated under vacuum, firing»
This process is perfromed before pump assembly and following internal outgassing process: You might probably know that Agilent has a unique «outgassing» process for all the manufactured pumps which consist in a automatic baking at 450°C base èressure dependant(10-8 mbar).

## Slide 18: Vacuum Firing Process
VIP200 baking process cycle
Standard VIP baking process cycle
Lower pressure achieved at the end of the baking procedure (approx. a decade lower)
Pressure Temperature
4 hrs
2hr30’’

## Slide 19: Summary
I/P values have been measured for a wide range of Ion Pumps key design parameters such as V , B , cell dimensions
The influence of B on I/P has been experimentally measured , and a value higher than the one used in existing ion pumps has been proved to be necessary to improve the performances, especially at lower pressures.
With optimized design parameters it is possible to design an ion pump with the highest pumping speed per unit volume ( 200 L/S in the same dimensions of conventional 150 L/S ).
Even more important, the maximum pumping speed of 200 L/S is reached in the 10 -8 mbar (10 -6 Pa ) range, where a conventional 150 L/S usually provides 120 L/S.
Vacuum firing process of all the surfaces exposed to vacuum allows to reach a lower ultimate pressure

## Slide 20
New Agilent VIP200
eXtreme High Vacuum Performance
The first ion pump with maximum pumping speed at low pressure