## Slide 1
Metrological characterisation of single-photon avalanche diodes
Marco López
Robin Eßling
Beatrice Rodiek
Helmuth Hofer
Stefan Kück
Working group 4.54: Laser radiometry and quantum radiometry

## Slide 2: Motivation
Single-Photon Avalanche Diodes (Si or InGaAs SPADs) important for
  scientific research:
    experimental quantum optics
    quantum cryptography
    quantum computing
  medicine
  biology
  astrophysics
   Wherever low photon fluxes need to be measured!

## Slide 3: Motivation- Light
Classical light: Super-Poissonian-Distribution  bunching (thermal light)
Laser light: Poissonian Distribution
Non-classical light: Sub-Poissonian-Distribution  anti bunching
  semiconductor quantum dot, single atoms / molecules / ions,
single organic molecules, single defect centres in crystals

## Slide 4: Motivation – Analog detector vs. Digital detector
Signal  Number of photons
Si
hn
Si
hn
I1
I2
I1 < I2
SPAD
hn
SPAD
hn
„Click“
„Click“
Analogue detector
Digital detector
Only one “click” per time

## Slide 5: Detector types
 | Si-SPAD | InGaAs-SPAD | TES | SNSPD
Detection efficiency | 80 % | 20 % | > 90 % | > 85 %
Dark counts | 5 cps | > 1 kHz < 50 Hz | - | < 10 cps
Jitter |  40 ps |  250 ps  100 ps |  100 ns | < 25 ps
Deadtime |  50 ns | > 1 µs | > µs | < 10 ns
Max. count rate |  20 MHz |  1 MHz | < 1 MHz | > 50 MHz
Photon number resolution | no | no | yes | (no)
After-pulsing |  0.5 % |  5 % | - | no
Spectral range | (350 – 1000) nm | (900 – 1700) nm |  – mm | (< 0.5 – > 2.5) µm
Operation temperature | RT | RT | LT | LT

## Slide 6: Detector types
 | Si-SPAD | InGaAs-SPAD | TES | SNSPD
Detection efficiency | 80 % | 20 % | > 90 % | > 85 %
Dark counts | 5 cps | > 1 kHz < 50 Hz | - | < 10 cps
Jitter |  40 ps |  250 ps  100 ps |  100 ns | < 25 ps
Deadtime |  50 ns | > 1 µs | > µs | < 10 ns
Max. count rate |  20 MHz |  1 MHz | < 1 MHz | > 50 MHz
Photon number resolution | no | no | yes | (no)
After-pulsing |  0.5 % |  5 % | - | no
Spectral range | (350 – 1000) nm | (900 – 1700) nm |  – mm | (< 0.5 – > 2.5) µm
Operation temperature | RT | RT | LT | LT
 | Si-SPAD | InGaAs-SPAD
Detection efficiency | 80 % | 20 %

## Slide 7: Detection efficiency, 
Photon detection probability (Detection efficiency, ): probability that a photon incident at the optical input will be detected within a detection gate.
Pi : Photon detection probability at each illuminated gate
Pd : Dark count probability
<n>: Average photon number per emitted pulse (mean photon number)
Average optical power [W]

## Slide 8: Measurement methods
Photon correlation technique (Reference detector not needed)
Calibrated
detector
SPAD
Substitution method (Reference detector needed)

## Slide 9: SPAD calibration using double attenuator technique – Step 1
M. López, et. al, “Detection efﬁciency calibration of single-photon silicon avalanche photodiodes traceable using double attenuator technique, Journal of Modern Optics 62, S21 – S27 (2015)
Laser 770 nm
Monitor detector
Filter 2
T = 4.6  10-4
Filter 3 T = 1.6  10-4
Si-Diode
Si-SPAD
Microscope objective
Beam splitter
Variable Filter
OD = 0.2 … 4

## Slide 10: SPAD calibration using double attenuator technique – Step 2
M. López, et. al, “Detection efﬁciency calibration of single-photon silicon avalanche photodiodes traceable using double attenuator technique, Journal of Modern Optics 62, S21 – S27 (2015)
Laser 770 nm
Monitor detector
Filter 2
T = 4.6  10-4
Filter 3 T = 1.6  10-4
Si-Diode
Si-SPAD
Microscope objective
Beam splitter
Variable Filter
OD = 0.2 … 4

## Slide 11: SPAD calibration using double attenuator technique – Step 3
M. López, et. al, “Detection efﬁciency calibration of single-photon silicon avalanche photodiodes traceable using double attenuator technique, Journal of Modern Optics 62, S21 – S27 (2015)
Laser 770 nm
Monitor detector
Filter 2
T = 4.6  10-4
Filter 3 T = 1.6  10-4
Si-Diode
Si-SPAD
Microscope objective
Beam splitter
Variable Filter
OD = 0.2 … 4

## Slide 12: SPAD calibration using double attenuator technique – Step 4
M. López, et. al, “Detection efﬁciency calibration of single-photon silicon avalanche photodiodes traceable using double attenuator technique, Journal of Modern Optics 62, S21 – S27 (2015)
Laser 770 nm
Monitor detector
Filter 2
T = 4.6  10-4
Filter 3 T = 1.6  10-4
Si-Diode
Si-SPAD
Microscope objective
Beam splitter
Variable Filter
OD = 0.2 … 4

## Slide 13: SPAD calibration using double attenuator technique – Step 5
Laser 770 nm
Monitor detector
Filter 2
T = 4.6  10-4
Filter 3 T = 1.6  10-4
Si-Diode
Si-SPAD
Microscope objective
Beam splitter
Variable Filter
OD = 0.2 … 4
M. López, et. al, “Detection efﬁciency calibration of single-photon silicon avalanche photodiodes traceable using double attenuator technique, Journal of Modern Optics 62, S21 – S27 (2015)

## Slide 14: Si-SPAD calibration using double attenuator technique
Laser 770 nm
Monitor detector
Filter 2
T = 4.6  10-4
Filter 3 T = 1.6  10-4
Si-Diode
Si-SPAD
Microscope objective
Beam splitter
Variable Filter
OD = 0.2 … 4
Si-SPAD
Si-Diode
F3 F2
Variable Filter
Monitor
Laser
Beam splitter
M. López, et. al, “Detection efﬁciency calibration of single-photon silicon avalanche photodiodes traceable using double attenuator technique, Journal of Modern Optics 62, S21 – S27 (2015)

## Slide 15: Measurement uncertainty budget – main components
Spectral responsivity of Si-diode:  40 %
  Caused by linearity measurements and its uncertainty
Filter transmission:  45 %
  Total filter transmission equals multiplication of two single filter transmissions?
SPAD = 0.6359  0.0040 (k = 2)
SPAD = 0.6359  0.63 % (k = 2)

## Slide 16: Measurement setup – Improvements
Standard detector:
  Integrating sphere with Si-diode instead of Si-diode only (avoiding back-reflexion into the setup)
Automated alignment procedure:
  Better reproducibility in position

## Slide 17: Detection efficiency – Results and Uncertainty
Uncertainty component | Uncertainty (%)
Planck constant, h | 2.52 x 10-7
Speed of light, c | 0.0
Wavelength, λ | 0.0075
Amplification factor, A1 | 0.0021
Amplification factor, A2 | 2.08 x 10-6
Amplification factor, A3 | 2.08 x 10-6
Ratio V1/VMon1, Q1 | 0.004
Ratio V2/VMon2, Q2 | 0.015
Ratio V3/VMon3, Q3 | 0.05
Ratio CR/VMonSPAD, Q4 | 0.036
Spectral responsivity, sSi | 0.15
Factor for the use of two filters, Ffilt | 0.005
Combined uncertainty, uc | 0.162
u(ηSPAD)  0.16 %
(770 nm, 100 000 cps)
Main contribution:
Standard detector
However, this is the ideal value. What will be achieved in day-to-day calibrations? 1 % seems reasonable!? Pilot study under way!!!

## Slide 18: (In)Homogeneity
Relative spatial detection efficiency of the Si-SPAD detectors
Laser beam diameter B approx. 10 µm.
Relative values are within ± 1 % for the main active region of the detectors, at the border of the active regions the values drop.
Homegeneities
Region 1: ≤ 0.85 % ( = 120 µm) Region 1: ≤ 2.2 % ( = 40 µm)
Region 2: ≤ 0.3 % ( = 40 µm) Region 2: ≤ 0.13 % ( = 20 µm)

## Slide 19: Detection efficiency – influence of photon statistics
M. López, et. al, “Detection efﬁciency calibration of single-photon silicon avalanche photodiodes traceable using double attenuator technique, Journal of Modern Optics 62, S21 – S27 (2015)

## Slide 20: InGaAs calibration
Reference detector:
Low-noise InGaAs photodiode (Hamamatsu G8605-23), cooled at -20 °C, and a Femto/Pico-ammeter (Keysight B2981A)
  Dark current: Id ~ 4 pA
  Linearity: < 0.25 % for Iph = 10 µA – 10 pA

## Slide 21: InGaAs calibration

## Slide 22: InGaAs calibration
Source of uncertainty | Standard uncertainty (%)
Planck´s constant, h | 9.055×10-6
Wavelength,  | 0.004
Speed of light, c | 0.000
Spectral Responsivity of InGaAs-diode, SInGaAs | 0.300
InGaAs/InP SPAD counts, Pcount | 0.720
InGaAs/InP SPAD dark counts, Pdark | 0.305
InGaAs/InP SPAD after pulsing, FPafter | 0.173
Photocurrent InGaAs diode, Iph | 0.028
Attenuation factor, FAtt | 0.050
Linearity factor InGaAs photodiode, FLin | 0.050
Combined uncertainty, uc | 0.857

## Slide 23: Traceability chart for cal. SPADs
Beam path:
1 – 2 – 3 – 2 – 1
 radiation nearly totally absorbed
Electrical Substitution Radiometer (ESR) operated at cryogenic low temperatures (4-7 K).
Optical power is traced back to electrical power (electric current I and electric potential difference V).
 = 0.999879
 = 0.99958
Sensitivity = 1.24 K/mW at 4.2 K

## Slide 24: (First) European pilot comparison: DE (free-running InGaAs-SPAD)
Device Under Test (DUT):
Free-running InGaAs SPAD (ID Quantique, ID 220)
: 1550 nm
Trigger Freq.: 110 kHz
U()= 2 % - 6 % (En  0.6)
SPIE Photonics, 22 - 26 April 2018, Strasbourg, France
If ,

## Slide 25: Summary
Calibration setup and procedure (Si and InGaAs SPADs)
Measurement uncertainty budget
  Standard measurement uncertainty < 0.8 %
Improvements (free-beam):
  Filter transmission measurements
  Standard measurement uncertainty 0.16 % (@Si SPAD)
  Homogeneity
Comparisons: Consistent results
  Double attenuator technique vs. CMI-Standard
  First European comparison (@InGaAs, 1550 nm)

## Slide 26: Acknowledgement
Thank you for your attention!