## Slide 1: Large scale characterization of SiPMs in the MEG II experiment
Y. Uchiyama (The University of Tokyo) on behalf of MEG II collaboration
International Conference on the Advancement of Silicon Photomultipliers 12 Jun, 2018 @ Schwetzingen

## Slide 2
Applications & usage of large number of SiPMs in MEG II
Not an R&D for future experiments but in use in a real experiment
  Detectors have already been constructed and operated.
  Mass pro. by companies, their quality, pre-test & characterization, calibration during the experiment, etc.
R&D
Construction/ commissioning
2012
2015
2019
Run
SiPM production in 2014–2015

## Slide 3: MEG II: in search of μ+ → e+γ
Search for μ+ → e+γ down to
6×10-14 (90% C.L. sensitivity)
B(μ+ → e+γ) < 4.2×10-13 (90% C.L.)
(while 5.3×10-13 expected)
×2 intensity muon beam
×2 resolution everywhere
×2 efficiency
μ+: World’s most intense DC muon beam @ PSI
γ : Detect with liquid xenon scintillation detector
e+: Detect with gradient B-field spectrometer (drift chamber & timing counter inside)
To get definitive evidence for BSM
An intensity frontier experiment
Upgraded from MEG experiment
    The upgrade largely relies on the application of SiPMs
MEG final result (2016)

## Slide 4
Liquid xenon photon detector (LXe)
Pixelated timing counter
(pTC)
Cylindrical drift chamber (CDCH)
COBRA
superconducting magnet
Radiative decay counter (RDC)
Muon stopping target
μ+
γ
e+
Scintillation detectors with SiPM
MEG II design (EPJ-C 78(2018)380)

## Slide 5: Two large-scale SiPM-based detectors
Pixelated Timing Counter
For e+ timing measurement
SiPM for the readout of plastic scint. counters.
Highly segmented design (512 counters)
12 SiPMs/counter, 6 (series connection)×2 sides
Total 6144 SiPMs
  From AdvanSiD
  3×3 mm2, 50 um pitch
Liquid Xenon photon detector
For γ energy, timing & position measurement
SiPM for the readout of liquid xenon scint. light
Highly granular readout (4092 ch)
Total 4092×4=16368 SiPMs
  From Hamamatsu Photonics
  Large size (12×12 mm2) by 4 chips in a package (‘hybrid’ connection)
  VUV sensitive (λ=175 nm)
  Used at low temperature (-100℃)
See W. Ootani’s talk (on Fri.) for the device and detector

## Slide 6: Readout electronics in pre-tests
During the R&D and mass test, we used standalone readout system.
KEITHLEY picoammeter
Amplifier developed at PSI (by U. Greuter)
DRS4 evaluation board (by S. Ritt)
  https://www.psi.ch/drs/evaluation-board
  4 ch waveform digitizer
  able to daisy-chain multiple boards
USB→PC

## Slide 7: Readout electronics in the experiment
New DAQ/Trigger system being developed: WaveDAQ system
  Used for all MEG-II detectors in common
  Dense & compact system to cope with increased # of channels. Away from VME crates
  No pre-amplifier at detector side
  Custom multi-functional readout board: WaveDREAM
    Analog FE (programmable shaper & amplifier), SiPM bias-voltage supply, waveform sampling (DRS4), digitization, discriminator, FPGA-based trigger in one module
  Synchronization accuracy < 20 ps (over different crate modules)
256 channels/crate
16 channels/module

## Slide 8: Readout electronics in the experiment
New DAQ/Trigger system being developed: WaveDAQ system
  Used for all MEG-II detectors in common
  Dense & compact system to cope with increased # of channels. Away from VME crates
  No pre-amplifier at detector side
  Custom multi-functional readout board: WaveDREAM
    Analog FE (programmable shaper & amplifier), SiPM bias-voltage supply, waveform sampling (DRS4), digitization, discriminator, FPGA-based trigger in one module
  Synchronization accuracy < 20 ps (over different crate modules)
256 channels/crate
16 channels/module
HV supply
2-stage amplifier
Pole-zero cancellation
Digitization
Trigger
Sync.
SiPM signal waveform digitized at 2.0 GSPS
baseline restoration by the pole-zero cancellation

## Slide 9
Liquid xenon photon detector

## Slide 10: LXe-MPPC structure
A VUV-sensitive large-area MPPC consists of 4 independent SiPM chips
  Chips with similar Vbd are selected.
  Able to readout individually.
  We connect them on the assembly PCB
A chip
6 mm
6 mm
An MPPC (S10943-4372)
An assembly PCB
‘Hybrid’ connection
a chip
~1 m2

## Slide 11: Mass test 1: individual chip test
To identify bad chips,
we measured I-V characteristics for all 4200×4 chips.
  Breakdown voltage (Vbd)
  Leakage current @ Vover = 5.0 V (Imeas)
  Shape of I-V curve
Temperature:20℃
~40 MPPCs (~160 chips) are measured at once.
relays to change the readout
chips
Setup for mass test 1
~250 MPPCs / day
Multiplexing PCBs

## Slide 12: Mass test 1: individual chip test
To identify bad chips,
we measured I-V characteristics for all 4200×4 chips.
  Breakdown voltage (Vbd)
  Leakage current @ Vover = 5.0 V (Imeas)
  Shape of I-V curve
compare with the data sheet (Vhama, Ihama)
Vhama – Vbd
Imeas – Ihama
Operation voltage
from data sheet
Measured breakdown voltage
Measured current at V = Vbd+5.0 V
Current at Vhama from data sheet
one outlier
large current chips
x ~10
Fraction of outlier chips is very small.
(gain=2.0x106 @ 25)

## Slide 13: Mass test 2: MPPC test on PCB
To test MPPC signals
Data taking with strong LED light
V=Vhama (operation voltage from spec. sheet, Vover~5V)
MPPC mounting + test = ~15min/PCB
186 PCBs  ~47hours in total
Waveform and charge were measured for all MPPCs on every PCB.
LED1
LED0
Temperature:20℃
reference MPPC

## Slide 14: Mass test 2: MPPC test on PCB
Vertical axis is charge normalized by charge of a MPPC in batch A.
Batch A
B
C
D
All channels are good.
Batch by batch difference is found, due to difference in afterpulse probability.
Batch A
B
C
D
Waveform of 4092 MPPCs overlaid.

## Slide 15: Signal check during construction
To check cabling and connection,
measure the MPPC currents with flashing LEDs in array one-by-one and check the correspondence.

## Slide 16: Signal check during construction
To check cabling and connection,
measure the MPPC currents with flashing LEDs in array one-by-one and check the correspondence.

## Slide 17: Close the detector & fill it with LXe
View from a USB camera installed inside the detector
LXe transfer

## Slide 18: MPPC calibration during RUN
The detector is equipped with calibration sources:
  LEDs
  Alpha sources (241Am) on wires
Gain
  Weak LED light(*)
      Measure 1 p.e. charge, with amp ×100
  Strong LED to monitor relative G
  Other methods under development
      Photon statistics at various LED intensities
      Waveform fluctuation
Crosstalk/afterpulse
  Weak LED, deviation from Poisson
PDE @ 175 nm
  Alpha signal
Time offset
  LED
241Am source
(*) DCR is too low at -100℃

## Slide 19
Charge distribution for weak LED light
1 p.e.
0 p.e.
Serial No.
all MPPCs
Gain change correlates with LXe temperature
PDE estimated from α
Uncertainty on the absolute value
  Light yield
  Integration range
Constantly monitor
A
B
C
D
all MPPCs
Serial No.
PDE for LXe light (175 nm)

## Slide 20: Large variation in afterpulse prob.
Batch dependence of afterpulse prob. was observed
Pulse shape difference, especially at high Vover.
Calibration & analysis become complicated
Vover = 7 V
Time offset calibration with LED
Measured signal timing
Distance correction
Batch (PCB) dependence is visible
Achieved calibration precision of ~60 ps
The correction depends on the pulse shape

## Slide 21
e+ pixelated Timing Counter

## Slide 22: Counter structure
Optical fiber for laser calib.
SiPM: ASD-NUV3S-P High-Gain (MEG)
  3×3 mm2, 50 μm pixel pitch
  Breakdown voltage: 24 V
  Operational range: Vover=2 – 3.5 V
  Peak sensitivity: 420 nm
  Temperature coeff. VBD: 26 mV/℃
  Production in 2014
A SiPM-array (a readout ch)
6 SiPMs connected in series
120 mm
50 mm
BC-422

## Slide 23: Mass test 1: individual SiPM test
To identify bad SiPMs and to group 6-SiPM sets,
we measured I-V characteristics for >7000 SiPMs.
  Breakdown voltage (Vbd)
  Leakage current @ Vover = 3.0 V (I3V)
Setup for mass test 1
Control board
Power suppply
Copper bench with electrical contact to SiPMs sitting on it controlled at 30℃
32 pcs at once
SiPM’s electrodes (surface mount type)
Spring probe pins for the contact
Group by the order of I3V

## Slide 24: Mass test 1: individual SiPM test
To identify bad SiPMs and to group 6-SiPM sets,
we measured I-V characteristics for >7000 SiPMs.
  Breakdown voltage (Vbd)
  Leakage current @ Vover = 3.0 V (I3V)
Group by the order of I3V
~7000 SiPMs overlaid
Narrow distribution of Vbd
We observed 2 groups in I-V curves
Type1-2 shows quick blowup of current
  Due to larger crosstalk/afterpulse at high Vover
  Result in narrower operation range → lower PDE

## Slide 25: Mass test 1: individual SiPM test
To identify bad SiPMs and to group 6-SiPM sets,
we measured I-V characteristics for >7000 SiPMs.
  Breakdown voltage (Vbd)
  Leakage current @ Vover = 3.0 V (I3V)
Group by the order of I3V
~7000 SiPMs overlaid
Narrow distribution of Vbd
We observed 2 groups in I-V curves
Type1-2 shows quick blowup of current
  Due to larger crosstalk/afterpulse at high Vover
  Result in narrower operation range → lower PDE

## Slide 26: Mass test 2: SiPM-array test
I-V measurement of 6-series SiPM array
Sensitivity to scintillation light (∝PDE)
Use the same scintillator for all
Exchange SiPM-array (coupled with grease)
Measure signal amplitude
@ room temperature
Setup for mass test 2 (simplified measurement)
Setup for detail test (time resolution measurement)

## Slide 27: Mass test 2: SiPM-array test
I-V measurement of 6-series SiPM array
Sensitivity to scintillation light (∝PDE)
Setup for mass test 2 (simplified measurement)
Setup for detail test (time resolution measurement)
Large variation of PDE observed
15% standard deviation
Bad ones has <1/2 PDE of typical (mean) one
Clear correlation b/w signal amplitude (∝PDE)
and time resolution

## Slide 28: Mass test 3: Counter test
Measure the time resolution of all the assembled counters
Find optimal operation point
Type 1-1
Type 1-2
Type 2
The variation includes scintillator quality
The difference b/w 40 & 50 mm counters is due to the sensor coverage
The optimal points depend on the SiPM types
Time resolution gets worse when the dark current starts to blowup
Type 1-2 has narrow optimal region and worse resolution.
Type 2 is a different batch (model) produced in 2015
Suppressed afterpulsing widen the operational range
But time resolution is not necessarily better

## Slide 29
Assembled & installed.
Sep. 2017

## Slide 30: Calibration during RUN
We are NOT interested in detailed parameters of SiPMs such as
  Gain, PDE, crosstalk/afterpulse prob.
But only in timing
  Timing & the time resolution
  Dark count rate (noise) ∝ dark current
The detector is equipped with
  Laser calibration system
  Temperature control & monitor system
  Continuous current monitor system
Laser light injected in a counter
(λ=401 nm, 60 ps FWHM)

## Slide 31: Current monitor
Currents are continuously monitored during the experiment at precision of a few nA
  Easily identify problems on SiPMs.
  When necessary, we can take I-V curves for all channels at once
Gradual increase caused by radiation damage
  →increase in the DCR
  Fluence is position dependent
  Need readjustment of operation voltage channel-by-channel time-to-time.
5 days
Beam ON
(μA)

## Slide 32: Calibration & optimization with laser
Calibrate time offset of each counter with the synchronous laser pulse signal.
Monitor the timing (time offset) with the laser pulse over time.
  also the pulse amplitude (∝G×PDE)
Optimize the bias voltage for the best timing resolution.
  under a real noise condition.
Laser power
7.5
9
RMS: 2.5 ps

## Slide 33: Conclusions
We established methods for mass test and calibration.
  Total inspection with a quick method is important before detector construction.
      reject bad sensors, optimize combinations/arrangement etc.
  Detailed characterization before experiment is important to understand the detector.
  Continuous monitoring of SiPM parameters is important during the experiment to maximize the performance.
      temperature change, radiation damage, etc.
Standardization makes little (no) sense once detector is built
  Necessity is different b/w detector R&D/designing and operation/calibration.
  We can/should only & all what we can do within the constraints of experiment.
Quality control in mass-production is important
  Not only to guarantee the high performance,
  but also to make the analysis & calibration easy.
  Hope manufacturers to improve this.

## Slide 35: Methods for Vbd
Linear + quadratic
2nd derivative of log(I)
@MPPC mass test
@LXe
@AdvanSiD SiPM mass test
The proper way depends on the environment and setup.

## Slide 36: PDE @175 nm
T.Doke et al., NIM A 420 (1999), 62.
Charge distribution of α events
Detailed simulation of photon tracking is necessary (reflection etc.)
Dependent on LXe purity
Notes: (\text{PDE})&=(\text{Measured p.e.})/(\text{Expected photons})\\
(\text{Measured p.e.})&=\frac{(\text{Measured charge})}{(\text{Gain})*(\text{Expected value of CTAP})}\\
(\text{Expected photons})&=\frac{(\text{Energy of Alpha source})}{19.6(\pm 2.0) \mathrm{eV} }*(\text{Solid Angle})

## Slide 37: Mass test in LXe
We did mass-test of ~600 pcs of prototype MPPCs in LXe
  Check the performance for LXe light →determine the final version
  Establish the procedure
MPPC &
support structure
feed-through
Xe gas panel
α sources
241Am
LED
feed-through
PCBs
(12+12)
24MPPCs / PCB
In LXe

## Slide 38: Radiation tolerance
Total fluence in MEG II
  ~1×1011 e+/cm2 in 3-year run.
  neutron flux is not severe
Impact on SiPM
  Dark current is expected to increase up to ~100 μA due to increased dark count rate.
  Tested impact on time resolution
    Using 37 MBq 90Sr source
    e+ beam at Beam Test Facility @ LNF
  30% degradation in time resolution @ 100 μA @ 30℃ 5% degradation @ 10℃
Before irradiation ~2 μA
~53 μA
~97 μA
~126 μA
~156 μA

## Slide 40: Series connection of SiPMs
SiPMs have high sensor capacitance
  Terminal capacitance ~300 pF for 3×3 mm2 SiPM
This forms a slow RC time-constant with amplifier input impedance
  3×3 mm2 : 300pF × 50Ω = 15 ns
  3×9 mm2 : 900pF × 50Ω = 45 ns !!
    → One of limitations for large area SiPMs or array of SiPMs with parallel connection
This large capacitance works as capacitive coupling when connected in series
SPICE simulation
Single cell
3×3mm2 MPPC
=
C
C
C
fired cells
the other cells + stray capacitance
~
Rshunt
e-
I
Measurement with 3 MPPCs