## Slide 1
June 11, 2018
HAMAMATSU PHOTONICS K.K
Solid State Division
Ardavan Ghassemi
Hamamatsu’s SiPM (MPPC) Characteristics and Latest Developments
1
1st Intl. Conference for Advancement of Silicon Photomultipliers
Schwetzingen, Germany
Notes: I’m belong to project promotion group and I’m working for promote the MPPC and MPPC module.
We will introduce the latest MPPC and module information.
もう一回ゆっくり話してもらえませんか？
Would you say that again more slowly, please ?
図に書いてもらえませんか？
Would you write that on the white board ?
いつでも質問してください
Please do not hesitate to ask questions during my presentation.

## Slide 2
2
Table of Contents
1. A Revision of MPPC Correlated Noise Parameters
2. Our Correlated Noise Measurement Method
3. New MPPC Developments for Physics Research

## Slide 3
3
Part 1
1. A Revision of MPPC Correlated Noise Parameters
2. Our Correlated Noise Measurement Method
3. New MPPC Developments for Physics Research

## Slide 4
4
Common questions from MPPC’s scientific users
How to distinguish delayed crosstalk from afterpulse?
How does Hamamatsu measure these noise parameters?
Why my measurements don’t match Hamamatsu’s?
What is a good method for comparing Hamamatsu’s MPPC with other manufacturers’ products?
Streamlined definitions of noise and other performance parameters can be helpful…

## Slide 5
<Main MPPC Meas. Parameters> ・ Photon detection efficiency ・ Dark count ・ Gain ・ Crosstalk ・ Afterpulese
PDE=detected photon count
/incident photon count
No problem, definition is clear.
Pulse count rate exceeding a threshold of 0.5p.e under dark condition.
No problem, definition is clear.
5
MPPC key performance parameters
Question : Which parameter is the toughest to define?
“Delayed” X-talk and afterpulse
are difficult parameters to fully distinguish…
Charge generated from one pixel when fired / Electron charge (1.602 x 10-19C).
No problem, definition is clear.

## Slide 6
6
MPPC correlated noise parameters
User Question: How to distinguish delayed crosstalk from afterpulse?
Hamamatsu Question: Is it really necessary to tell them apart?
Overvoltage = 3 V
50 ns
S13360-1350CS
Which one is afterpulse?
Which is delayed crosstalk?

## Slide 7
7
Revision of MPPC correlated noise parameters
No practical need to differentiate between delayed crosstalk and after-pulse!
Let’s redefine correlated noise parameters so that noise measurement becomes simpler and also more objective.
Correlated noise can be fully categorized into 2 parameters based on just one metric: time delay!
Hamamatsu’s proposal for new correlated noise parameters:
Prompt crosstalk & “Delayed pulses”

## Slide 8
8
Definition of prompt crosstalk
1 : Prompt crosstalk
( secondary photons reaching a neighboring cell and triggering a secondary avalanche )
※１Reference
(1a) Direct X-talk
(1b) Front-surface reflected X-talk
(1c) Bottom-surface reflected X-talk
These 3 are
“Prompt X-talk”

## Slide 9
9
Waveform example – prompt crosstalk
Measured pulse
Overvoltage = 6 V
Prompt X-talk
50 ns
1p.e.
2p.e.
Prompt X-talk
≧ 2p.e.
Shaped pulse

## Slide 10
2 : Delayed pulses
( minority carriers generated in the substrate diffuse to the depletion region with time delay, triggering a secondary avalanche )
Diffuse to the primary cell, pulse height depends on cell’s recovery progress.
10
Definition of delayed pulses
※１Reference
(2a) Afterpulses
(2b) Delayed X-talk
Diffuse to a neighboring cell,
pulse height is same level as 1p.e.
These 2 are
“Delayed pulses”

## Slide 11
11
Waveform examples – delayed pulses
Measured pulses
Shaped pulses
Delayed X-talk
= 1p.e.
Afterpulses
≦ 1p.e.
Overvoltage = 6 V
Delayed X-talk
Afterpulses
50 ns
Delayed pulses
1p.e.
Note: Delayed X-talk and afterpulses are indistinguishable after full recovery of microcell voltage.

## Slide 12
12
Part 2
1. A Revision of MPPC Correlated Noise Parameters
2. Our Correlated Noise Measurement Method
3. New MPPC Developments for Physics Research

## Slide 13
13
Measuring MPPC characteristics
User questions from an earlier slide:
Why my measurements don’t match Hamamatsu’s?
How does Hamamatsu measure correlated noise?
What is the best method to compare Hamamatsu’s MPPC with other suppliers’ products?
Replicating the same measurement method reproduces the same result, so we share our method with you!

## Slide 14
Focus on 1pixel
Correlated noise measurement
・ Configuration block

## Slide 15
・ Instrumentation setup
Overall view
PC
Black box
PLP
Controller
PLP
Head
Bias supply
for MPPC
Oscilloscope
Correlated noise measurement

## Slide 16
・ Biasing and readout scheme:
Correlated noise measurement
Gp：20.5dB
fc：2.9GHz
Gp：20.5dB
fc：1.0GHz
-HV
MPPC
Output
Amplifier
Amplifier
Capacitor
Capacitor
Capacitor
Capacitor
Capacitor
Resistor
Resistor
Resistor
Inductor
+12V

## Slide 17
Problem : Signal Pileup
In the waveform shown below, pulse height is affected by pileup.
Fig.5
Amplifier output
Pileup
Pulse processing

## Slide 18
How to alleviate pileup : software processing
The raw signal is passed through Deconvolution and
Wiener filters to obtain a processed output signal.
Filter
×
Pulse processing
Pileup is removed.

## Slide 19
Data points: ~10,000 (exp.)
MPPC
AMP(×100)
Oscilloscope
Computer
PLP
Normal
Prompt X-talk
Delayed Pulse
Filtering & Discrimination
0
0
0
2
3
1
1
1
Correlated noise measurement

## Slide 20
Waveform analysis ① : Prompt X-talk probability
It is defined by the following calculation:
Prompt X-talk(%)＝
(number of >1.5 p.e. pulses)
(number of >0.5 p.e. pulses)
×100
Correlated noise measurement
1.5pe
0.5pe
Delayed X-talk
After pulse
Signal

## Slide 21
Waveform analysis ② : Delayed Pulse probability
It is defined by the following calculation:
Delayed Pulse(%)＝
(number of Delayed pulses)
(number of >0.5 p.e. pulses)
×100
Signal
Signal
After pulse
Delayed X-talk
0.5pe
0.5pe
Correlated noise measurement

## Slide 22
22
Part 3
1. A Revision of MPPC Correlated Noise Parameters
2. Our Correlated Noise Measurement Method
3. New MPPC Developments for Physics Research

## Slide 23
23
Improved irradiation lifetime
2 types of damage caused by radiation in MPPC:
Bulk damage: If sufficient particle energy is transferred to a Si atom, the atom can be displaced into an interstitial site, leaving a vacancy (defect) in the Si crystal lattice.
Surface damage: Energetic photons and electrons can cause surface damage near the Si-passivation (SiO2/Si3N4) boundary.
Both can result in increased dark output vs. accumulated irradiation.
Proposed Solution:
New MPPC with lower dark current
W/ high dynamic range (small microcells) for calorimetry

## Slide 24
24
High DR w/ reduced dark output
Small microcell with trench(HDR2)
  Low bias operation (~45V)
  High PDE
  Low Dark Current
Under Development

## Slide 25
25
MPPC for Cherenkov detection
S14520* Series
*Mainly developed for CTA

## Slide 26
26
VUV-MPPC
S13370 series
High PDE in VUV wavelength range
Cryogenically reliable
Low optical crosstalk due to trenches
Fast recovery (stable Rq vs. temperature)
*Array type and low-RI pckg under development.

## Slide 27
27
MPPC connection technique
3 connection methods are adopted in multichannel applications.
Series and hybrid connection suitable for cryogenic or high-count-rate applications, requiring reduced capacitance.
Technical note is being prepared...
Series
Hybrid
Parallel
Small bias voltage
Large capacitance
(long waveform)
High bias voltage
(Vop x 4)
Small capacitance
(short waveform)
Bias : parallel
Signal: series
(coupled with C)
Small bias voltage
Small capacitance

## Slide 28
28