Progress of Lifetime measurement of MCP-PMT BINP #82
Alexander Britting, Wolfgang Eyrich, Albert Lehmann, Fred Uhlig
sponsored by BmBF and GSI

---
Setup of the Illumination and Gain measurement
Photodiode for monitoring of LED stability
2. Scaler for reduction to 'monitoring mode'

---
Quantum efficiency measurement
wavelength selection by grid
anode and PMT are shorted and current measured

---
How do we measure?
Measurement of the ADC-spectra (1 Scaler active)
Illumination, Measurement of the ADC-spectra in 'monitoring mode' = reduced event rate (few Hz) with both Scalers active
2. Measurement of the ADC-spectra (1 Scaler active)
QE-Measurement

---
Gain and Number of PE
Gain remains constant, PE decreases by factor ~5
Blackout destroyed +60V channel of HV, resulting in Gain 'gap'

---
Darkcurrent
Measurement error for 5mC/cm²
strong darkcurrent while 'blackout-measurements'

---
spectral QE
decrease of QE with higher Illumination
for low wavelengths: darkcurrents dominates because of spectral distribution of the light source

---
relative QE
QE drops significantly faster for higher wavelengths

---
QE(3)
QE drops by 62%(400nm) to 360%(600nm)
decrease slower than PE in ADC spectra

---
possible explanation for difference QE/PE
ions drift back to the protection layer => thickness of layer increases
increasing thickness of the protection layer results in a decrease of the probability for electrons passing the layer => CE decreases
QE measurement unchanged
ADC-spectra
QE-measurement

---
CE measurement
CE drop should be measureable by varying Cathode-MCP-Voltage
Higher Voltage needed to collect all electrons

---
Illuminationtime
Chargeconsumption decreases with increasing time, because of decreasing PE
jump at 4,65*10⁶ s: increase of the illumination rate (500kHz instead of 200kHz)

---
Outlook
Measurement of CE drop
For further lifetime measurements a higher accuracy should be achievable by:
    - better stabilised positioning for QE-measurement
Lifetime measurement for Photonis 85012 are currently being prepared