## Slide 1: HEBT Magnets Batch2&3 Magnetic Measurements
07.11.2018
C. Mühle

## Slide 2: Magnetic measurements at GSI
First results of magnetic measurement at GSI
  1st pair of S100 steerers delivered in August 2018
    Hall probe mapping done
      Complete set-up with girder
      Complete set-up with girder but ordinary steel bushes/screws between magnet and holders replaced by SS ones (S100v)
      Magnet w/o girder as at BINP (S100h)
  1st quadrupole 2 delivered November 5th 2018
    Planned
      Mapping with gradient hall probe
      Rotating coil measurement
      Axis measurements with alignment cylinder (hall probe)

## Slide 3: Prototype quadrupole 2
Delivered on November 5th 2018
No damages
Documentation not complete yet

## Slide 4: S100 steerer pair on girder

## Slide 5: S100 steerer coil movement
Video
Temporary repair
Improved design (still ~0.5mm movement

## Slide 6: Comparison with and w/o girder 1
∫Bdl=f(x)
I=400A
With girder: black-y=0, pink-y=-65mm, blue-y=65mm
W/o girder: red-y=0, green-y=-65mm, yellow-y=65mm

## Slide 7: Comparison with and w/o girder 2
∫Bdl/∫Bdl(x=0)=f(x)
Not x=0, y=0.
I=400A
With girder: black-y=0, pink-y=-65mm, blue-y=65mm
W/o girder: red-y=0, green-y=-65mm, yellow-y=65mm

## Slide 8: Comparison with and w/o girder 3
W/o girder: black-y=0, pink-y=-65mm, blue-y=65mm
With girder: red-y=0, green-y=-65mm, yellow-y=65mm
∫Bdl/∫Bdl(x=0)=f(x)
I=400A

## Slide 9: Comparison bushes/screws ordinary steel/stainless steel 1
∫Bdl=f(I)
x=0 y=0
Linear fit (black curve)
subtracted
GSI: Ordinary steel: black-up, pink-down
Stainless steel: blue-up, red-down
BINP: Ordinary steel: green

## Slide 10: Comparison bushes/screws ordinary steel/stainless steel 2
B=f(I)
x=0 y=0
Linear fit (black curve)
subtracted
GSI: Ordinary steel: black-up, pink-down
Stainless steel: blue-up, red-down
BINP: Ordinary steel: green

## Slide 11: Comparison BINP/GSI 1
BINP: blue-w/o girder
GSI: black–with girder , pink-w/o girder
∫Bdl=f(I)
x=0 y=0
Linear fit (black curve)
subtracted
Gain between BINP and
GSI ~1.01%

## Slide 12: Comparison BINP/GSI 2
BINP: blue-w/o girder
GSI: black–with girder , pink-w/o girder
∫Bdl=f(I)
x=0, y=0
Linear fit (individual)
subtracted
Gain between BINP and
GSI subtracted

## Slide 13: Comparison BINP/GSI 3
B(in center)=f(I)
x=0, y=0
Linear fit (black curve)
subtracted
Gain between BINP and
GSI ~1.01%
BINP: blue-w/o girder
GSI: black–with girder , pink-w/o girder

## Slide 14: Comparison BINP/GSI 4
B(in center)=f(I)
x=0, y=0
Linear fit (individual)
subtracted
Gain between BINP and
GSI subtracted
BINP: blue-w/o girder
GSI: black–with girder , pink-w/o girder

## Slide 15: Conclusions
Visual check during powering up and under power -> Will be part of FAT for prototypes in future
Magnetic measurement in completely assembled position mandatory
Improved coil fixation for S100
Improvements for girder of S100
Absolute magnetic field values to be checked