## Slide 1
Update on lifetime measurements
of MCP-PMTs and results of Photonis 9002150
D. Miehling, M. Böhm, A. Lehmann,
M. Pfaffinger, S. Stelter
PANDA-Meeting, June 25, 2019

## Slide 2: The ANDA detector at FAIR/GSI
Barrel DIRC
Disc DIRC
Focal planes of both
DIRC detectors are inside magnetic field
High annihilation rate
(average 20 MHz)
Notes: Wegen das Magnetfeldes von bis zu 2 Tesla MCP-PMT nötig

## Slide 3: The PANDA DIRC detector
Sensor requirements:
  Immunity to magnetic fields > 1 Tesla
  High rate stability (> 1 MHz/cm²)
  Fast time response (< 50 ps)
  Low dark counts (< 10 kHz/cm²)
  High position resolution:
  Compact design
  Microchannel-plate Photomultipliers (MCP-PMTs) needed
    Problem until few years ago: low lifetime
Lifetime indicator: Quantum efficiency as function of integrated anode charge (IAC) (detected number of photons)
    Barrel DIRC: 5x5mm²
    Disc DIRC: 0.5x16mm²
  Barrel DIRC IAC ~ 5 C/cm²
  Disc DIRC IAC > 5 C/cm²

## Slide 4: Aging of photocathode
Photocathode of MCP-PMT gets damaged from impact of (heavy) feedback ions produced in the residual gas
Damaging effects at the photocathode
  Chemical reactions and adsorption
  Surface defects
Possible countermeasures against aging
  More robust photocathode
  Better vacuum
  Reducing flux of (heavy) ions

## Slide 5: Approaches to increase Lifetime
New cathode material
  Baking in Cs/Sb –vapor
    Disadvantage: higher dark count rate
Protection layer (film
  In front of first MCP layer
    Disadvantage: reduction of collection efficiency
  Between MCP layers
Treatment of MCP surfaces
  Electron scrubbing
  Atomic layer deposition (ALD)
    Company Arradiance & LAPPD
    First prototypes from Photonis
    Later prototypes from Hamamatsu, too
Best results and focus of this talk:

## Slide 6: Parameters of lifetime measurements
Goal: Simultaneous measurement of different MCP-PMTs under similar conditions as in the PANDA-DIRCs
Constant illumination with 1 MHz single photons
  All MCP-PMTs in same light spot
  Permanent monitoring of integrated anode charge
QE measurement:
  Every few weeks:
    Wavelength spectrum with Xenon arc lamp (75 W) and a monochromator (Δλ = 2 nm, 250 nm - 800 nm)
    Also measuring gain and dark counts
  Every several months:
    Surface scans with picosecond laser (372 nm, spot size: ø~0.5-1 mm)

## Slide 7: Lifetime setup
Pulsed LED inside aluminum Box
Thorlabs engineered diffuser
  In front of LED to get homogeneous light spot on sensors
Holding construction for up to 16 sensors all illuminated by same LED
Cable management behind sensors
LED box with diffuser
Voltage dividers
Sensors
Cable holder
Lens

## Slide 8: Illumination Overview QE (all sensors with ALD)
Manufacturer
Senor ID
IAC
[mC/cm²]
QE latest [%]
QE latest/QE start [%]
QE start [%]
Photonis
XP85112
Hamamatsu R10754X
Hamamatsu R13266-07-M768 / M64
9001223
9001332
9001393
22.1
5.3
24
15909
23.0
8.2
36
19.1
19.3
101
KT0001 (M16M)
19534
21.7
5.6
26
KT0002 (M16M)
18175
21.1
6.2
JS0022 (64 pix.)
608
17.4
4.9
28
9234
97
18
JS0035 (64 pix.)
5294
25.5
24.8
JS0018 (768 pix.)
1133
18.0
JS0027 (768 pix.)
2261
24.3
22.2
91
3.22
17942
29
Film between MCP
Two ALD layers
2 Inch
2 Inch
1 Inch
Film in front of first MCP
Data from February 21, 2018

## Slide 9: Illumination Overview QE (all sensors with ALD)
Manufacturer
Senor ID
IAC
[mC/cm²]
QE latest [%]
QE latest/QE start [%]
QE start [%]
Photonis
XP85112
Hamamatsu R13266-07-M768 / M64
9001393
9002108
23938
19.1
19.6
103
21.7
10.7
49
100
96
JS0035 (64 pix.)
9911
25.5
25.4
JS0027 (768 pix.)
4182
24.3
YH0250 (64 pix.)
6151
25.4
25.2
99
23.3
1534
Two ALD layers
2 Inch
2 Inch
Film in front of first MCP
Data from June 24, 2019

## Slide 10: QE scans of new Hamamatsu 2 inch ALD devices
JS0027 (6x128 pix., film)
JS0035 (8x8 pix., film)
Covered
later prototypes tend to show better performance
measurement still ongoing
YH0250 (8x8 pix., no film)
JS0022 (8x8 pix., film)

## Slide 11: QE scans of Photonis 9001393-URD (2 ALD layers)
Not covered
2 ALD layers
No damage visible @24 C/cm² (~48 years of PANDA)

## Slide 12: QE scans of Photonis 9002108 (ALD)
Probably caused by feedback ions
ALD-coating by Photonis (and not Arradiance) 
Aging starting from the center
Not illuminated side less and later damaged
Covered (not illuminated)
Clear sign of Cathode damage

## Slide 13: Lifetime data of all sensors (June 24, 2019)
Most sensors with ALD coated MCPs have lifetime > 5 C/cm²
devices with no countermeasures have a lifetime of < 200mC/cm²
Measurements are taking a long time
No countermeasures
Film
New cathode material
ALD

## Slide 14: Summary
Requirements: > 5 C/cm² at gain (50% duty cycle, 10 years)
lifetime increased by a factor of 50-100 with ALD coating
Photonis
  Best sensor at 24C/cm² without any sign of cathode damage
  New sensor already damaged at 1.2C/cm²
Hamamatsu:
  Later produced (higher serial number) 2 inch tubes tend to have better performance
  JS0035 reached 5 C/cm² as first 2 inch tube (now at almost 10C/cm²)

## Slide 15: Results of Photonis 9002150 (Part 1)

## Slide 16: Gain vs voltage
Measured with scope
Signal not amplified
Gain calculated with gauss fit in histogram data (pedestal and signal)
gain at 2250V (with 250kΩ:1MΩ:125kΩ divider!) ~ 1650V MCP voltage
datasheet: 1725V MCP voltage
charge/gain spectrum
and signals
at 2250V

## Slide 17: QE vs wavelength
-200 V at cathode, ground at MCP-in
Current measured at MCP-in
Calculated in reference to photo diode current (known QE for each wavelength)
Data below 280nm and above 650nm not very reliable  light source
Max QE ~20% at 400nm

## Slide 18: QE surface scan
-200 V at cathode
Current measured at MCP-in
Calculated in reference to photo diode current (known QE for wavelength)
Scanned with 372 nm (blue)
0.5 mm steps across surface

## Slide 19: Current gain scan
Scanning 2200 V (~ gain)
Measuring shortened anode current
Scan data are folded with QE of the sensor
  Have to be divided by QE
QE corrected gain shown in picture on the right side
Gain then scaled to known value of one pixel
Area of
known
gain
x
Artificial edge due to few hours of interruption of scan (darkcurrent settled in this time)

## Slide 20: Time resolution
Measured with scope at 2250V (~ gain)
Red PiLas at 1 kHz and 27 % tune
200x amplified signal, followed by impedance matched splitting
Low discriminator threshold (30mV) (just above noise band)
Time walk corrected spectra
= 31 ps (RMS = 185 ps)
log bild bis 4ns

## Slide 21: Rate stability
Measured with shortened anodes at gain
Full illuminated sensor and photodiode in same light spot (diffuser)
Monitor diode and shortened anode current for different laser frequencies
In theory, when diode current doubles anode current should double too
Norm to first value taken:
should be 1 if sensor doesn’t change its gain
Better than 9002085 and 9002108
Higher rates were not accessible because of limitation of anode current

## Slide 22: Photonis 9002150
explanation of ringing slides
9002150
ND2.3
6*0.7=
4.2 ph
x: 10ns/div
y: 10mV/div
Trigger:
Laser pulse
signal of three different pixels (with persistence, ~5k events):
C1: pix26 (yellow)
C2: pix44 (red)
C3: pix76 (blue)
charge spectrum (histogram) + log of charge spectrum (thin line) of all events(~100k)
C1: pix26 (yellow+blue)
C2: pix44 (red+green)
C3: pix76 (missing+red)
colour of arrow  signal colour
oval  charge spectrum
log of charge spectrum  filling
number of illuminated pixels * number of photoelectrons (determined with the charge spectrum histogram and poisson statistics) = number of photons seen by the sensor

## Slide 23: Photonis 9002108 and 9002150
2 inch x 2 inch, 8 x 8 pixels, 10 µm pores with ALD coating, gain
Red (632 nm) PiLas, 10 kHz, illumination of 6 pixels
9002150
ND2.3
6*0.7=
4.2 ph
9002108
ND1.6
6*0.7=
4.2 ph
x: 10ns/div
y: 10mV/div
Trigger:
Laser pulse
here: C1=pix26 (crosstalk pixel)
here: C1=pix16 (normal pixel)

## Slide 24: Photonis 9002108 and 9002150
2 inch x 2 inch, 8 x 8 pixels, 10 µm pores with ALD coating, gain
Red (632 nm) PiLas, 10 kHz, illumination of 12 pixels
9002150
ND2.0
12*2=
24 ph
9002108
ND1.3
12*2=
24 ph
x: 10ns/div
y: 10mV/div
Trigger:
Laser pulse

## Slide 25: Photonis 9002108 and 9002150
2 inch x 2 inch, 8 x 8 pixels, 10 µm pores with ALD coating, gain
Red (632 nm) PiLas, 10 kHz, illumination of all pixels
9002150
ND2.3
64*7=
450 ph
9002108
ND1.6
64*7=
450 ph
x: 10ns/div
y: 10mV/div
Trigger:
Laser pulse

## Slide 26
Thank you for your attention!

## Slide 27: New feature: afterglowing?
ratestability measurement with shortenend anodes and current measurement
turning laser on and off again with increasing light intensity
with monitoring of light intensity  can determine ratestability
anode current in A
time in s

## Slide 28: New feature: afterglowing?
left: Photonis 9001393-URD
right: Hamamatsu YH0250
red line corresponds to 2MHz photonrate per cm²
anode current in A
time in s
anode current in A
time in s

## Slide 29: New feature: afterglowing?
this feature seems to occur with both Photonis and Hamamatsu tubes and affects only ALD-tubes
no effect: Photonis 9001341 (non-ALD tube)
almost no or a small effect: Photonis 9001393 (two-ALD), 9001394 (ALD), 9002108 (ALD), 9002150 (ALD); Hamamatsu JS0026 (ALD?, Gießen), JS0035 (ALD)
a strong effect:
Hamamatsu YH0245 (ALD?, Gießen), YH0250 (ALD)

## Slide 30: New feature: afterglowing?
comment: the summary on the last slide is preliminary since (most of) the data is from old ratestability measurements and only „bycatch“
more investigations needed with both current measurements and PADIWA and/or DiRiCH