## Slide 1
Photonis 2150, DiRICH and Photonis ES440
Merlin Böhm, A. Lehmann, S. Krauss, D. Miehling,

## Slide 2: Measurement setups and measured parameters
Surface scans: PILAS Laser with micro focus attached to a 3-axis stepper
FPGA based DAQ: TRB, Padiwa
  Padiwa FEE for discrimination
  TRB for time and TOT measurement
  Multihit capability
TRB3 and Padiwa3:
  Signal height distribution (important for setting the threshold)
  Crosstalk, charge cloud width and recoil electron behavior
  For each Pixel: time resolution, darkcount rate and afterpulse probability
y
x
z
Photodiode
MCP-PMT
Laser with ND-filter
or diffusor
y
x

## Slide 3: Photonis 9002150 - TRB measurements
Narrow time window of 1.5 ns, at 2200V (0.7e6 gain)
2 Pixels simultaneously hit reveals charge cloud width
Charge cloud width ~1.6 mm FWHM
3 Pixels simultaneously hit reveals edge of pixels
1 Hit
2 Hits
3 Hits

## Slide 4: Photonis 9002150 - TRB measurements
Narrow time window of 1.5 ns, at 2200V (0.7e6 gain)
2 Pixels simultaneously hit reveals charge cloud width
Charge cloud width ~1.6 mm FWHM
3 Pixels simultaneously hit reveals edge of pixels
1 Hit
2 Hits
3 Hits
Two channels disabled
Shorted

## Slide 5: Photonis 9002150 - TRB measurements
Shorted pixels are visible in 2 hit plots
Pixels are shorted inside the MCP-PMT
1 Hit
2 Hits
3 Hits

## Slide 6: TRB vs DiRICH
Irradiating one pixel and alter the threshold
Compare 1pF and 10pF at DiRICH3 frontend lowpass filter and Padiwa3

## Slide 7: TRB vs DiRICH
Idea: use same threshold for comparison
Create plots with Laser/Pixel time difference vs threshold
Padiwa3 (with modified capacitor and resistor):

## Slide 8
DiRICH 1 pF:
DiRICH 10 pF:
Time difference
Time over threshold

## Slide 9: TRB vs DiRICH
About 10-20% improvement in time resolution between 10 pF (~125 ps RMS) compared to 1 pF (~140 ps RMS) and Padiwa 3
Need to check higher capacitor values
DiRICH 10 pF
Padiwa3

## Slide 10: Photonis ES440
3x100 Pixel MCP-PMT
Narrow time window of 1.5 ns at 1e6 gain
1 Hit:
2 Hits:
QE:

## Slide 11: Photonis ES440
Time resolution (5 ns window): ~140 ps
Sensor darkcount rate: ~20 kHz
Average afterpulse probability: ~0.25 %

## Slide 12: Photonis ES440
Further investigating the time distribution
Trigger cut (TC): only events when the laser is on the pixel

## Slide 13: Photonis ES440 – time distributions
gh

## Slide 14: Photonis ES440 – Hit distributions
f

## Slide 15
Plot mean of the fits of the individual pixels
Pixel position is linear (fit)

## Slide 16
Plot width (sigma) of the fits of the individual pixels
Pixel width mostly constant

## Slide 17: Summary
Photonis 9002150 has two shorted pixels
Higher low pass capacitance at DiRICH frontend leads to lower signal heights but better time resolutions
Photonis ES440: first measurements look promising, good time resolution, low afterpulse probability and good position accuracy